THE PROVISION OF AUTOMATIC EXTERNAL DEFIBRILLATORS (AEDs), MINI AED TRAINING UNITS AND ASSOCIATED CONSUMABLES

MINIMUM TECHNICAL SPECIFICATIONS

Category - General.

- The device must be easily transportable in its case with all accessories.

- The device must have a water ingress protection (IP) rating of 4 or above (as defined in accordance with BS EN 60529:1992 or equivalent).

- The device must be able to be stored at 0 - 40 degrees C.

- The device must be able to operate at 0 - 40 degrees C.

- The device must be able to operate at non-condensing relative humidity levels of 10 - 95%.

- The device must incorporate a realtime clock and/or the associated software must provide the ability for events to be time-stamped.

- The device must have a life expectancy of at least 5 years.

- Replaceable batteries and/or any other replaceable power sources for the device must have a shelf life of at least 3 years.

- Replaceable batteries and/or any other replaceable power sources for the device must have a life expectancy of at least 2 years once inserted into the defibrillator.

- The device must carry out analysis during CPR with no need to pause. There should be no requirement for delays in CPR for heart analysis to take place

Category - Power.

- The device must be able to operate on a battery only basis.

- The device must have a battery shock capacity minimum of 30 full discharges at the maximum energy your device operates at throughout the standard life of the battery.

- The initial analysis time plus the time to charge and be ready to deliver initial shock must be a maximum of 15 seconds.

Category - Warranty.

- The device must have a warranty of at least 4 years.

- The battery and/or any other replaceable power source for the device must have a warranty of at least 3 years.

Category - Data / Memory.

- The device must have the capability to store data on at least one of the following: PC Data Card or internal memory.

Category - Facilities.

- The device must have simple controls which can be understood by a non-expert.

- The device must have voice prompt ability in English including a safety message.

- The device must have visual prompt ability.

- The device must have biphasic waveform or a waveform of comparable efficacy.

- The device must have automatic discharge control to safely discharge unit if charge not used (i.e. when device charged but patient not shocked).

- The device must be able to shock VF and VT above 150 BPM.

- The device must have impedance control.

- The device must have a low battery level alarm or low battery level indicator.

- The device must have an indicator confirming self-test success - i.e. indicates safe to use.

- The device must be able to withstand daily cleaning with a detergent wash and disinfection using a range of nationally available generic disinfection agents without any detrimental effect.

- There must be a carry case available for the device and related accessories.

- The device must be compatible with the most recent Resuscitation Council (UK) and / or European Resuscitation Council Guidelines.

- Average time from notification of unit error to repair, replacement or provision of a loan unit/consumables must be no more than five working days.

- The device must not require any chargeable maintenance during its standard warranty period.

Category - Accessories.

- The shelf life of the self-adhesive pads for this device for use on patients must be a minimum of 24 months from date of supply.

- The pads must have the ability to be pre-connected.

- The purchase price must include (as a minimum) one set of pads

- Outdoor cabinets must have a minimum IP rating of IP65.