## Equipment Specification
| User Requirement Specification | CPI Tender Reference & Title:- | 2245 | Supply & Delivery of Accelerated Weathering Tester
Lot # and Title:-
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column C, of which the total Quality element score of this sheet will equate to value in cell F6 | Scoring Criteria 0 = no response/answer not relevant to requirement 1 = limited response to requirements 2 = less than adequate, filfils some of the basic requirements 3 = adequate response, fulfils all basic requirements 4 = good response, fulfils all requirements to a high standard 5 = excellent response, fulfils requirements to an exceptionally high standard |  | The total percentage available which will contribute towards the Quality scoring element of this tender | 50%
USER REQUIREMENT |  |  |  | TENDERER RESPONSE
Requirement Codes |  |  |  | Guidance
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as desirable) |  |  | Tenderers must provide responses within the following columns; 1. Column F - please provide details of your equipment functionality in relation to CPI requirements specified in column E 2. Column G - please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response |  | Y
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  |  | N
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
# | Parameter | Requirement | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N) | N
System Performance
M1 | Light source | Broadband, high intensity light source capable of emitting visible and ultra-violet wavelengths. | M |  |  | 1 | 0 | 0
M2 |  | Light intensity must be automatically controlled with a feedback circuit such as an irradiance sensor. | M
M3 |  | Ability to filter/tailor the light source wavelengths, by using interchangable filters, for example. | M
D1 |  | Light source can be filtered to accurately mimic the spectral output of direct sunlight. | D
O1 |  | Light source can be filtered to extend the spectral output into the ultra-violet regions. | O
D2 |  | Irradiance values of at least 70 W/sq. m at 300-400 nm achievable for >24 hours using daylight filters | D
D3 |  | Light source classified as "ozone free" | D
M4 | Sample chamber environmental control | Chamber air temperature can be set to any temperature between 25-65 ºC during the light cycle. | M |  |  | 4 | 1 | 3
M5 |  | Chamber air temperature can be set any temperature between 15-50 ºC during dark cycle. | M
D4 |  | Channel air temperature can be maintained to within ± 2 ºC of the programmed value | D
D5 |  | Chamber relative humidity can be specified between 30-80% rH (± 5% rH) at 25 ºC, when external conditions are 21 ºC, 50 % rH. | D
M6 |  | Chamber air temperature sensor installed. | M
D6 |  | Insulated black panel temperature sensor installed. | D
D7 |  | Simultaneous control of chamber air tempertaure and black panel temperature. | D
M7 | Specimen compatibility and chamber capacity | The total available test area for specimens must be 2.8 m² or greater. | M |  |  | 7 | 1 | 7
M8 |  | Test area can be configured to hold coated substrates, such as alumninium or steel panels | M |  |  | 8 | 1 | 7
D8 |  | Test area can be configured to hold mechanical testing dog bones: ISO527 type 5 (115 mm long, 25 mm wide, > 1 mm thick) and ISO527-2-1BA (75 mm long, 10 mm wide and > 5 mm thick) | D |  |  | 8 | 1 | 8
D9 |  | Test area can be configured to hold 3-dimensional objects up to 10 cm in height. | D
D10 |  | Test area can be configured to hold combinations of different specimen types (panels/dog bones/3D objects). | D
M9 | Water spray | Ability to spray samples with clean water to simulate rain. | M
M1 |  | Water recycling system to minimise water consumption/wastage. | M
D11 |  | Monitoring of water quality with total dissolved solids sensor | D
| Installation and Utilities |  |  |  |  | 10 | 1 | 11
M11 | Installation | Installation must be performed by qualifed engineers at CPI NetPark facility | M |  |  | 11 | 1 | 11
M12 | Site Requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment. | M |  |  | 12 | 1 | 11
M13 | Electrical Power | The system should operate from UK mains supply (240Vac, 50Hz, single phase or three phase) | M |  |  | 13 | 1 | 11
D12 | Other Utilities | Options to upgrade with additional capabilities for future proofing | D |  |  | 13 | 1 | 12
| Software and Connectivity |  |  |  |  | 13 | 1 | 12
D13 | Connectivity | Test cycles are programmable without the need of an external computer e.g. via touch screen input, or similar. | D |  |  | 13 | 1 | 13
D14 |  | Data logging over ethernet or USB. | D
| Regulatory, Support and Warranties |  |  |  |  | 13 | 1 | 14
M14 | CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided. | M |  |  | 14 | 1 | 14
M15 | Training | Full on-site training provided along with access to technical support. | M |  |  | 15 | 1 | 14
D15 | Warranties | Warranty period should cover at least 12 months from installation | D |  |  | 15 | 1 | 15
D16 | Extended Warranty | Extended warranty of up to 10 years | D |  |  | 15 | 1 | 16
M16 | Service Support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. | M |  |  | 16 | 1 | 16
M17 |  | Details on required methods for regular calibration to be provided | M
M18 |  | Details of preventative maintenance programmes required (if any) above Warranty requirements | M |  |  | 18 | 1 | 16
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com