## Drop Codes (hidden)
RoomClassification | IA | QualificationTest
Not Classified | Direct Impact | Yes
Grade A | Indirect Impact | No
Grade B | No Impact | n/a
Grade C
Grade D

## Appendix 1
USER REQUIREMENT - SPECTROMETER
Requirement codes
S | Mandatory Safety Requirement (Bio, toxicity, flammable, manual handling etc)
M | Mandatory
O | Optional requirment which may be quoted for and an assessment made by CPI prior to purchase
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  | Answer | Requirement (Q/S/M/D/N) -see codes | Response
System performance
| Instrument principles | The system should allow online powder compostion analysis using vibrational spectroscopy. | M
|  | The system should allow the determination of material (e.g., carboxymethyl cellulose, sodium bicarbonate and citric acid) composition at line or off-line during the progression stages of pharmaceutical and food processing operations. | M
Technical specifications
| Thermal stability studies | The system should allow spectral acquisition at a pre-defined temperature setting that falls within the range from 0ºC and 100ºC | D
| Operating temperature range | The system should ideally operate in a temperature range of 5°C - 30°C | D
| Humidity range | The system should ideally operate in a humidity range of 20 – 80% | D
| Sample format | The system should allow the analysis of both solid samples | D
| Sample volume and concentration | For offline measurments sample volume/quantity requirements should be provided as part of tender repsonse. | D
Installation & Utilities
| Installation | Installation must be performed by qualifed engineers at CPI Formualtion facility | S
| Qualification | Software and hardware IQ/OQ packages must be provided for the appropriate equipment | D
| Site requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment | M
| Electrical Power | The system should operate from single phase (13A or below) supply | D
Software and Connectivity
| Connectivity | Ethernet TCP/IP connection | D
| Instrument Software & Hardware | System specific software, fittings & PC should be supplied with the instrument | D
| Integration | Capabilities for integration to laboratory informatics systems should be detailed | D
| Software | Specific software should be included to provide analysis of protein secondary structure (for example, helix, sheet and other content) | D
| Software | The software should support instrument connectivity via TCP/IP, setup and operation, data collection, data processing, and export to file formats such as .csv , .xls and PDF. | D
Regulatory, Support & Warranties
| FDA CFR 21 Part 211 Part D - Equipment | The bidder should provide a statement of compliance | D
| FDA CFR 21 Part 1040 Performance Standards for Light-Emiting Product | The bidder should provide a statement of compliance | D
| FDA 21 CFR Part 11 | The bidder should provide a statement of compliance | D
| CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided | M
| Training | Full hardware and software training programmes should be provided along with access to application specific technical support | D
| Warranties | Extended Warranty period should cover 36 months from installation | D
| Service support | Details of preventative maintenance programmes required (if any) above Warranty requirements | D

## Appendix 2
USER REQUIREMENT - PARTICLE SIZE ANALYSER
Requirement codes
S | Mandatory Safety Requirement (Bio, toxicity, flammable, manual handling etc)
M | Mandatory
O | Optional requirment which may be quoted for and an assessment made by CPI prior to purchase
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  | Answer | Requirement (Q/S/M/D/N) -see codes | Response
System performance
| Instrument principles | A particle analyser, for online analysis of dry powders and bulk solids to proivide particle size and shape information for batch and continuous processes. The instrument should be capable of analysing a wide range of formulated materials from a range of sectors including pharmaceuticals, FMCG, personal care, chemicals and foods. | M
Technical specifications
| Size range | Should be able to measure particle size from 1 micron to 3000 micron (Additional capability to measure particle sizes above 5000micron is desirable) | M
| Mode of analysis | Online, at-line and off-line | M
| Operating temperature range | Measure in processes with temperature range of 5°C to 50°C | D
Installation & Utilities
| Installation | Installation must be performed by qualifed engineers at CPI Formualtion facility | S
| Qualification | Software and hardware IQ/OQ packages must be provided for the appropriate equipment | D
| Site requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment | M
| Electrical Power | The system should operate from single phase (13A or below) supply | D
Software and Connectivity
| Connectivity | Ethernet TCP/IP connection | D
| Instrument Software & Hardware | System specific software, fittings & PC should be supplied with the instrument | D
| Integration | Capabilities for integration to laboratory informatics systems should be detailed | D
| Software | Specific software should be included to provide analysis of protein secondary structure (for example, helix, sheet and other content) | D
| Software | The software should support instrument connectivity via TCP/IP, setup and operation, data collection, data processing, and export to file formats such as .csv , .xls and PDF. | D
Regulatory, Support & Warranties
| FDA CFR 21 Part 211 Part D - Equipment | The bidder should provide a statement of compliance | D
| FDA CFR 21 Part 1040 Performance Standards for Light-Emiting Product | The bidder should provide a statement of compliance | D
| FDA 21 CFR Part 11 | The bidder should provide a statement of compliance | D
| CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided | M
| Training | Full hardware and software training programmes should be provided along with access to application specific technical support | D
| Warranties | Extended Warranty period should cover 36 months from installation | D
| Service support | Details of preventative maintenance programmes required (if any) above Warranty requirements | D

## Appendix 3
USER REQUIREMENT - MOISTURE ANALYSER
Requirement codes
S | Mandatory Safety Requirement (Bio, toxicity, flammable, manual handling etc)
M | Mandatory
O | Optional requirment which may be quoted for and an assessment made by CPI prior to purchase
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  | Answer | Requirement (Q/S/M/D/N) -see codes | Response
System performance
| Instrument principles | The system should allow online moisture measurement in powders. | M
|  | The system should allow the determination of moisture in chemical powders, food powders, salt, dried food, fermented food, snack food and sugar at-line or off-line during the pharmaceutical and food processing upstream or downstream process stages | M
Technical specifications
| Moisture measurement range | 0-95% | M
| Maximum permissible absolute error of moisture measurement | +/- 0.2% | M
| Reproducibility | +/- 0.01% | D
| Working temperature range | 5°C - 50°C | D
Installation & Utilities
| Installation | Installation must be performed by qualifed engineers at CPI Formualtion facility | S
| Qualification | Software and hardware IQ/OQ packages must be provided for the appropriate equipment | D
| Site requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment | M
| Electrical Power | The system should operate from single phase (13A or below) supply | D
Software and Connectivity
| Connectivity | Ethernet TCP/IP connection | D
| Instrument Software & Hardware | System specific software, fittings & PC should be supplied with the instrument | D
| Integration | Capabilities for integration to laboratory informatics systems should be detailed | D
| Software | Specific software should be included to provide analysis of protein secondary structure (for example, helix, sheet and other content) | D
| Software | The software should support instrument connectivity via TCP/IP, setup and operation, data collection, data processing, and export to file formats such as .csv , .xls and PDF. | D
Regulatory, Support & Warranties
| FDA CFR 21 Part 211 Part D - Equipment | The bidder should provide a statement of compliance | D
| FDA CFR 21 Part 1040 Performance Standards for Light-Emiting Product | The bidder should provide a statement of compliance | D
| FDA 21 CFR Part 11 | The bidder should provide a statement of compliance | D
| CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided | M
| Training | Full hardware and software training programmes should be provided along with access to application specific technical support | D
| Warranties | Extended Warranty period should cover 36 months from installation | D
| Service support | Details of preventative maintenance programmes required (if any) above Warranty requirements | D

## Appendix 4
USER REQUIREMENT - BULK DENSITY MEASUREMENT
Requirement codes
S | Mandatory Safety Requirement (Bio, toxicity, flammable, manual handling etc)
M | Mandatory
O | Optional requirment which may be quoted for and an assessment made by CPI prior to purchase
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  | Answer | Requirement (Q/S/M/D/N) -see codes | Response
System performance
| Instrument principles | The system should allow online measurement of bulk density of powders. | M
|  | The system should allow the determination of bulk density in chemical powders, food powders, salt, dried food, fermented food, snack food and sugar at-line or off-line during the pharmaceutical and food processing upstream or downstream process stages | M
Technical specifications
| Measurement range | 0 kg/m3 to 2500 kg/m3 | M
| Measurement accuracy | +/- 0.2% | M
| Reproducibility | +/- 0.01 kg/m3 | D
| Operating temperatue | 5°C - 40°C | D
| Vibration limit | 4 Hz to 800 Hz | D
| Diamter of pipe | 20 mm to 100 mm | D
Installation & Utilities
| Installation | Installation must be performed by qualifed engineers at CPI Formualtion facility | S
| Qualification | Software and hardware IQ/OQ packages must be provided for the appropriate equipment | D
| Site requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment | M
| Electrical Power | The system should operate from single phase supply (12-24VDC) | D
Software and Connectivity
| Connectivity | Ethernet TCP/IP connection | D
| Instrument Software & Hardware | System specific software, fittings & PC should be supplied with the instrument | D
| Integration | Capabilities for integration to laboratory/pilot plant informatics systems should be detailed | D
| Software | Specific software should be included to provide analysis of protein secondary structure (for example, helix, sheet and other content) | D
| Software | The software should support instrument connectivity via TCP/IP, setup and operation, data collection, data processing, and export to file formats such as .csv , .xls and PDF. | D
Regulatory, Support & Warranties
| FDA CFR 21 Part 211 Part D - Equipment | The bidder should provide a statement of compliance | D
| FDA CFR 21 Part 1040 Performance Standards for Light-Emiting Product | The bidder should provide a statement of compliance | D
| FDA 21 CFR Part 11 | The bidder should provide a statement of compliance | D
| CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided | M
| Training | Full hardware and software training programmes should be provided along with access to application specific technical support | D
| Warranties | Extended Warranty period should cover 36 months from installation | D
| Service support | Details of preventative maintenance programmes required (if any) above Warranty requirements | D