## Drop Codes (hidden)
RoomClassification | IA | QualificationTest
Not Classified | Direct Impact | Yes
Grade A | Indirect Impact | No
Grade B | No Impact | n/a
Grade C
Grade D

## Autoclave
User Requirement Specification |  |  | CPI Internal Tender Reference & Title:- | 1073 - Autoclave and Upgrades
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 30%
USER REQUIREMENT |  |  |  |  | Tenderer Response
Requirement codes |  |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as Desirable) |  |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  | Requirement |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
| Equipment principle | General purpose electrically heated laboratory autoclave, suitable for sterilising equipment up to 1m in height. Programmable with process timer and USB connection for data logging and external communication. |  | M
| Application | Utilised for sterilisation solid and liquid loads e.g. lab scale fermentation equipment, growth media, shake flasks and contaminated waste cycles |  | M
Technical Sp[ecification
| Operating temperature | Able to achieve a constant temperature of at least 121°C. Unit must be capable of load sensed timing by inclusion of a temperature probe within the autoclave chamber. This is sometimes referred to as a 'Wandering Probe'. |  | M
| Dimensions | Internal height of at least 1m |  | M
| Control | Programmable temperature and time settings, with measurement of actual achieved temperature and data logging. |  | M
| Operation | End of cycle cooling by suitable means e.g. cooling fan |  | M
| Operation | Vacuum assisted drying |  | D
| Operation | Automatic water level control with auto-refill |  | D
| Safety | Built in Safety protection from excessive temperature and pressure |  | M
| Safety | Built in Safety protection to prevent opening door under high temperature or pressure. |  | M
| Hoist | supply of an appropriate hoist to lift heavier equipment into and out of the autoclave. |  | D
| Data logging | Built in printer to provide a hard copy of the temperature, pressue and time profiles for each autoclave cycle. |  | M
Installation & Utilities
| Installation | Installation must be performed by qualifed vendor engineers at CPI facility located at Wilton Centre |  | M
| Site requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment if required |  | M
| Utilities | All bids must include a full list of requried utilities |  | M
Regulatory, Support & Warranties
| CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided |  | M
| Training | Full training should be provided along with access to technical support |  | D
| Warranties | Warranty period should cover at least 18 months from installation |  | M
| Extended warranty | Extended warranty of up to 5 years |  | D
| Service support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. |  | D
| Service support | Details of preventative maintenance programmes required (if any) above Warranty requirements |  | D
OFFICE USE ONLY - This section below this line is for CPI Evaluation and Internal USE only
| Notes for Evaluation Team: |  | Name of Bidder to be inserted here | 1073 - Autoclave and Upgrades | 0
Please assess all of the answers given, with particular attention to the Mandatory answers provided, taking into account, the Desirable and Optional answers too, however, you do not need to score each individual answer, but based on the scoring range 0-5, just give one singluar score (default set at "5" in the red and yellow box below, and as we have to provide feedback to bidders, you must also provide comments on which section(s) above where you felt that they did not meet the requirements.
| Evaluation Panel |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com