## A - 4 port glovebox
| User Requirement Specification | CPI Tender Reference & Title:- | 1636 | Supply Delivery Installation Comissioning Argon Glovebox system
|  | Lot # and Title:- | Lot 1 - Option A | 1-4 Port Glovebox
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column C, of which the total Quality element score of this sheet will equate to value in cell F6 | Scoring Criteria 0 = no response/answer not relevant to requirement 1 = limited response to requirements 2 = less than adequate, filfils some of the basic requirements 3 = adequate response, fulfils all basic requirements 4 = good response, fulfils all requirements to a high standard 5 = excellent response, fulfils requirements to an exceptionally high standard |  | The total percentage available which will contribute towards the Quality scoring element of this tender | 50%
USER REQUIREMENT |  |  |  | TENDERER RESPONSE
Requirement Codes |  |  |  | Guidance
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as desirable) |  |  | Tenderers must provide responses within the following columns; 1. Column F - please provide details of your equipment functionality in relation to CPI requirements specified in column E 2. Column G - please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response |  | Y
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  |  | N
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
# | Parameter | Requirement | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N) | N
System Performance
M1 | Key instrument capability | 4-ports glovebox . Glovebox operating under Argon gas. | M |  |  | 1 | 0 | 0
M2 |  | Stainless steel construction | M |  |  | 2 | 0 | 0
M3 |  | Polycarbonate window for glovebox front. | M |  |  | 3 | 0 | 0
M4 |  | Closed loop recirculation | M |  |  | 4 | 0 | 0
M5 |  | Touch display panel for operating glovebox | M |  |  | 5 | 0 | 0
M6 |  | Positive pressure operation | M |  |  | 6 | 0 | 0
M7 |  | Footpedal for adjustment of system pressure | M |  |  | 7 | 0 | 0
M8 |  | Purifier system capable of maintaining atmosphere with minimum leakage. | M |  |  | 8 | 0 | 0
D1 |  | Ability to operate at reduced power or gas rates to save energy. | D |  |  | 8 | 0 | 1
M9 | Construction of system | 4-ports glovebox - single sided. | M |  |  | 9 | 0 | 1
M10 |  | Internal workspace dimensions: At least 1900 mm (W) x 750 mm (D) x 900 mm (H) | M |  |  | 10 | 0 | 1
M11 |  | Left hand large antechamber: at least 380 mm (Diameter) x 580 mm (Length) | M |  |  | 11 | 0 | 1
M12 |  | Left hand mini antechamber: at least 150 mm (Diameter), 400 mm (Length) | M |  |  | 12 | 0 | 1
|  | Above left hand mini antechamber fitted with a heating jacket capable of heating to temperatures >120 °C. | O
M13 |  | All antechambers fitted with sliding tray. | M |  |  | 13 | 1 | 1
M14 | Purging of antechambers | All antechambers able to be evacuated using vacuum and refilled with Argon | M |  |  | 14 | 1 | 1
D2 |  | Interlocks present on all antechambers to prevent exposure of internal system to external atmosphere. | D |  |  | 14 | 1 | 2
O2 |  | Ability to argon purge (with time contol) large antechamber without vacuum. | O |  |  | 14 | 2 | 2
M15 |  | Supplied with all necessary oil pumps for generation of vacuum. | M |  |  | 15 | 2 | 2
D3 |  | Capability to automatically evacuate and refill large antechambers with a couple of cycles. | D |  |  | 15 | 2 | 3
M16 | Moisture and oxygen monitoring | Capability to maintain environment with < 2 ppm O2 and moisture levels | M |  |  | 16 | 2 | 3
M17 |  | Solid state analyser to indicate when system rising above these levels and system regen required | M |  |  | 17 | 2 | 3
M18 | Solvent trapping | Molecular sieve based trapping system for collection of solvents. Must be suitable for working with solvent volumes > 200 mL. | M |  |  | 18 | 2 | 3
M19 |  | Solvent system able to be regenerated independently using control panel to remove solvent from trap automatically | M |  |  | 19 | 2 | 3
M20 | Feed throughs | Multiple feedthroughs along the length of glovebox to allow power cables to be routed | M |  |  | 20 | 2 | 3
M21 |  | At least 4 13A plugs supplied | M |  |  | 21 | 2 | 3
M22 |  | At least 4 blank (KF) flanges for future capability | M |  |  | 22 | 2 | 3
D4 |  | Ability to feed in up to 4 sets of wires from a potentiostat system (3 electrode configuration) | D |  |  | 22 | 2 | 4
M23 | Regeneration of catalyst system | Glovebox catalyst/reactor to allow regeneration using H2/Argon gas mixture. | M |  |  | 23 | 2 | 4
D5 |  | Automatic regeneration routine which switches circulation back on once completed, and stops reform gas supply | D |  |  | 23 | 2 | 5
M24 | Shelves/racking etc | Shelving supplied at rear of each glovebox for storage of materials | M |  |  | 24 | 2 | 5
O3 |  | Ability to slide shelves towards operator to enable easy access | O |  |  | 24 | 3 | 5
M25 | Spares/Consumables etc | Standard gloves of butyl rubber, at least 0.4 mm thick | M |  |  | 25 | 3 | 5
O4 |  | Please provide pricing detail for gloves of other materials | O |  |  | 25 | 4 | 5
M26 |  | Spare gloves provided for all ports. | M |  |  | 26 | 4 | 5
M27 |  | Include port inserts for changing gloves | M |  |  | 27 | 4 | 5
0 | Installation and Utilities |  |  |  |  | 27 | 4 | 5
M28 | Installation | Installation must be performed by qualifed engineers at CPI NetPark facility | M |  |  | 28 | 4 | 5
M29 | Site Requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment | M |  |  | 29 | 4 | 5
D6 | Electrical Power | Where possible, the system should operate from a standard UK mains supply (240Vac, 50Hz, single phase). Please indicate if an alternate supply is required. | D |  |  | 29 | 4 | 6
0 | Software and Connectivity |  |  |  |  | 29 | 4 | 6
D7 | Connectivity | Please indicate capability to data log eg for monitoring moisture, O2 levels. | D |  |  | 29 | 4 | 7
0 | Regulatory, Support and Warranties |  |  |  |  | 29 | 4 | 7
M30 | CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided | M |  |  | 30 | 4 | 7
M31 | Training | Full training should be provided along with access to technical support. Flexibility to give on-site training throughout the life of the equipment | M |  |  | 31 | 4 | 7
M32 | Warranties | Warranty period should cover at least 18 months from installation | M |  |  | 32 | 4 | 7
M33 | Extended Warranty | Extended warranty of up to 5 years | M |  |  | 33 | 4 | 7
M34 | Service Support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. | M |  |  | 34 | 4 | 7
M35 |  | Details of preventative maintenance programmes required (if any) above Warranty requirements | M |  |  | 35 | 4 | 7
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com

## B - Combined 4+3 glovebox
| User Requirement Specification | CPI Tender Reference & Title:- | 1636 | Supply Delivery Installation Comissioning Argon Glovebox system
|  | Lot # and Title:- | Lot 1 - Option B | Combined 4 + 3 Glovebox
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column C, of which the total Quality element score of this sheet will equate to value in cell F6 | Scoring Criteria 0 = no response/answer not relevant to requirement 1 = limited response to requirements 2 = less than adequate, filfils some of the basic requirements 3 = adequate response, fulfils all basic requirements 4 = good response, fulfils all requirements to a high standard 5 = excellent response, fulfils requirements to an exceptionally high standard |  | The total percentage available which will contribute towards the Quality scoring element of this tender | 50%
USER REQUIREMENT |  |  |  | TENDERER RESPONSE
Requirement Codes |  |  |  | Guidance
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as desirable) |  |  | Tenderers must provide responses within the following columns; 1. Column F - please provide details of your equipment functionality in relation to CPI requirements specified in column E 2. Column G - please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response |  | Y
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  |  | N
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
# | Parameter | Requirement | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N) | N
System Performance
M1 | Key instrument capability | 4-ports glovebox attached to a 3-ports glovebox. Glovebox operating under Argon gas. Entire assembly not more than 6 m in length | M |  |  | 1 | 0 | 0
M2 |  | Stainless steel construction | M |  |  | 2 | 0 | 0
M3 |  | Polycarbonate window for glovebox front. | M |  |  | 3 | 0 | 0
M4 |  | Closed loop recirculation | M |  |  | 4 | 0 | 0
D1 |  | Separate touch display panels for operating both gloveboxes independently | D |  |  | 4 | 0 | 1
M5 |  | Positive pressure operation | M |  |  | 5 | 0 | 1
M6 |  | Footpedal for adjustment of system pressure for each system | M |  |  | 6 | 0 | 1
M7 |  | Single purifier system capable of maintaining atmosphere in both boxes simultaneously, with minimum leakage. | M |  |  | 7 | 0 | 1
D2 |  | Ability to operate at reduced power or gas rates to save energy. | D |  |  | 7 | 0 | 2
M8 | Construction of 4-ports system | 4-ports glovebox - single sided. | M |  |  | 8 | 0 | 2
M9 |  | Internal workspace dimensions: At least 1900 mm (W) x 750 mm (D) x 900 mm (H) | M |  |  | 9 | 0 | 2
M10 |  | Left hand large antechamber: at least 380 mm (Diameter) x 580 mm (Length) | M |  |  | 10 | 0 | 2
M11 |  | Left hand mini antechamber: at least 150 mm (Diameter), 400 mm (Length) | M |  |  | 11 | 0 | 2
O1 |  | Above left hand mini antechamber fitted with a heating jacket capable of heating to temperatures >120 °C. | O |  |  | 11 | 1 | 2
M12 |  | All antechambers fitted with sliding trays. | M |  |  | 12 | 1 | 2
M13 | Interconnection | T shape interconnecting large antechamber: at least 380 mm (Diameter) x 580 mm (Length) with a sliding tray and an exterior door | M |  |  | 13 | 1 | 2
M14 | Construction of 3-ports system | 3 port glovebox - single sided. | M |  |  | 14 | 1 | 2
D3 |  | Internal workspace dimensions: At least 1500 mm (W) x 750 mm (D) x 900 mm (H) | D |  |  | 14 | 1 | 3
M15 |  | Right hand mini antechamber: at least 150 mm (Diameter), 400 mm (Length) | M |  |  | 15 | 1 | 3
M16 |  | Antechamber fitted with sliding tray | M |  |  | 16 | 1 | 3
D4 |  | Ability to install larger equipment through an accessible back panel at the right rear of the 3 port glovebox (eg SalvisLab Vacucenter VC20 Vacuum Oven, 60 x 40 x 45 cm in size) | D |  |  | 16 | 1 | 4
M17 | Purging of antechambers | All antechambers able to be evacuated using vacuum and refilled with Argon | M |  |  | 17 | 1 | 4
D5 |  | Interlocks present on all antechambers to prevent exposure of internal system to external atmosphere. | D |  |  | 17 | 1 | 5
O2 |  | Ability to argon purge (with time contol) large antechamber on 4 port glovebox, and interconnecting antechamber, without vacuum. | O |  |  | 17 | 2 | 5
M18 |  | Supplied with all necessary oil pumps for generation of vacuum. | M |  |  | 18 | 2 | 5
D6 |  | Capability to automatically evacuate and refill large antechambers with a couple of cycles. | D |  |  | 18 | 2 | 6
M19 | Moisture and oxygen monitoring | Capability to maintain environment with < 2 ppm O2 and moisture levels | M |  |  | 19 | 2 | 6
M20 |  | Solid state analyser to indicate when system rising above these levels and system regen required | M |  |  | 20 | 2 | 6
M21 | Solvent trapping | Molecular sieve based trapping system for collection of solvents. Must be suitable for working with solvent volumes > 200 mL. | M |  |  | 21 | 2 | 6
M22 |  | Solvent system able to be regenerated independently using control panel to remove solvent from trap automatically | M |  |  | 22 | 2 | 6
M23 | Feed throughs | Multiple feedthroughs along the length of glovebox to allow power cables to be routed | M |  |  | 23 | 2 | 6
M24 |  | At least 4 13A plugs supplied | M |  |  | 24 | 2 | 6
M25 |  | At least 4 blank (KF) flanges for future capability | M |  |  | 25 | 2 | 6
D7 |  | Ability to feed in up to 4 sets of wires from a potentiostat system (3 electrode configuration) | D |  |  | 25 | 2 | 7
M26 | Regeneration of catalyst system | Glovebox catalyst/reactor to allow regeneration using H2/Argon gas mixture. | M |  |  | 26 | 2 | 7
D8 |  | Automatic regeneration routine which switches circulation back on once completed, and stops reform gas supply | D |  |  | 26 | 2 | 8
M27 | Shelves/racking etc | Shelving supplied at rear of each glovebox for storage of materials | M |  |  | 27 | 2 | 8
O3 |  | Ability to slide shelves towards operator to enable easy access | O |  |  | 27 | 3 | 8
M28 | Spares/Consumables etc | Standard gloves of butyl rubber, at least 0.4 mm thick | M |  |  | 28 | 3 | 8
O4 |  | Please provide pricing detail for gloves of other materials | O |  |  | 28 | 4 | 8
M29 |  | Spare gloves provided for all ports. | M |  |  | 29 | 4 | 8
M30 |  | Include port inserts for changing gloves | M |  |  | 30 | 4 | 8
0 | Installation and Utilities |  |  |  |  | 30 | 4 | 8
M31 | Installation | Installation must be performed by qualifed engineers at CPI NetPark facility | M |  |  | 31 | 4 | 8
M32 | Site Requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment | M |  |  | 32 | 4 | 8
D9 | Electrical Power | Where possible, the system should operate from a standard UK mains supply (240Vac, 50Hz, single phase). Please indicate if an alternate supply is required. | D |  |  | 32 | 4 | 9
0 | Software and Connectivity |  |  |  |  | 32 | 4 | 9
D10 | Connectivity | Please indicate capability to data log eg for monitoring moisture, O2 levels. | D |  |  | 32 | 4 | 10
0 | Regulatory, Support and Warranties |  |  |  |  | 32 | 4 | 10
M33 | CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided | M |  |  | 33 | 4 | 10
M34 | Training | Full training should be provided along with access to technical support. Flexibility to give on-site training throughout the life of the equipment | M |  |  | 34 | 4 | 10
M35 | Warranties | Warranty period should cover at least 18 months from installation | M |  |  | 35 | 4 | 10
M36 | Extended Warranty | Extended warranty of up to 5 years | M |  |  | 36 | 4 | 10
M37 | Service Support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. | M |  |  | 37 | 4 | 10
M38 |  | Details of preventative maintenance programmes required (if any) above Warranty requirements | M |  |  | 38 | 4 | 10
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com