## Drop Codes (hidden)
RoomClassification | IA | QualificationTest
Not Classified | Direct Impact | Yes
Grade A | Indirect Impact | No
Grade B | No Impact | n/a
Grade C
Grade D

## URS Document
User Requirement Specification |  |  | CPI Internal Tender Reference & Title:- | 1187 | Framework for the Supply of Chromatography Systems
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 30%
USER REQUIREMENT |  |  |  |  | Tenderer Response
Requirement codes |  |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as Desirable) |  |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  | Requirement |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
1 | System Specification | Pumps must be compatible with high salt buffers |  | M
2 |  | Flow rate range of 0.5 - 100mL/min |  | M
3 |  | Must have 4 buffer feed per input channel |  | M
4 |  | The system must be able to run gradients |  | M
5 |  | pH monitoring is required |  | M
6 |  | Conductivity monitoring is required |  | M
7 |  | Pressure, air and temperature monitoring and control alarms are required |  | M
8 |  | On line buffer preparation |  | D
9 | UV Monitoring | Please detail wavelength range covered and the number of wavelengths that can be simultaneously recorded during operation, As a minimum, system must be able to record two wavelengths simultaneously. |  | M
10 |  | Post Column monitoring |  | M
11 |  | Flow cell diameter suitable for flow rate range without creating excessive back pressure. |  | M
12 |  | UV monitor with a long lifetime and automatic calibration |  | M
13 | Fraction Collector | System must be able to control fraction collection including programmable volumes or time based collection and a range of different triggers to fraction collection such as peak detection. |  | M
14 |  | Integrated fraction collector with ability to use different tube sizes. |  | M
15 |  | Temperature control of fractions is desirable |  | D
16 | Scalability | The system functionality should cover the ability to rapidly scale processes from screening and method optimisation using small columns, to fine-tuning and robustness testing using larger columns. Please provide details on the ability to use the system to develop scalable chromatography methods and processes |  | M
17 |  | System control methodology and hardware range must be scalable to demonstrate chromatography can be performed at pilot (100-200L) and manufacturing (100-1000L) scales, including GMP manufacture. To demonstrate scalability, please provide product details for larger scale equipment that is available as options |  | M
18 | Columns | System should be suitable for using columns of approximately 10-500ml packed bed volume |  | M
19 |  | System should be able to use pre-packed and self packed columns |  | M
20 |  | If columns are available as options, please provide details and quotes |  | M
21 |  | Ability to use multiple columns is necessary. Provide details |  | M
22 |  | Ability to perform both manual column packing, please provide details of columns/accessories required and quotes |  | M
23 |  | Ability to perform automated packing, please provide details of columns/accessories required and quotes |  | D
24 | Control System/Software | System must provide a suitable control system, specifically designed to run chromatography systems. The system must record all measured data and enable flexible process control. |  | M
25 |  | User friendly interface that can be easily configured by users |  | M
26 |  | Control software that allows both programmable and real-time control of the chromatography system |  | M
27 |  | User friendly interface that enables easy writing, viewing and editing of the chromatography method and run properties | User friendly interface that enables easy writing, viewing and editing of the chromatography method and run properties | M
28 |  | System must provide the option to export data for inclusion into reports |  | M
29 |  | Design of experiments functionality is desirable |  | D
30 | Design | Innovative design & features that give benefits such as flexibility and ease-of-use |  | M
31 | Dimensions | Please provide details of the system footprint |  | M
32 | Integration | Please detail approach to system integration such as network options and connection to a laboratory management system (LMS). |  | D
Installation & Utilities
33 | Installation | Installation must be performed by qualified engineers at CPI Darlington facility |  | M
34 | Site requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment |  | M
35 | Electrical Power | The system should operate from a standard UK mains supply (240Vac, 50Hz, single phase) |  | M
36 | Other Utilities | The system will not require a 3 phase supply, cooling water, gases or cryogenic gases |  | D
Software and Connectivity
37 | Connectivity | Computer Interface (e.g. ethernet, USB etc) |  | D
38 | Software | Software should be included for method development and data analysis |  | M
39 | Computer | Computer or laptop to run the software needed. |  | M
Regulatory, Support & Warranties
| CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided |  | M
| Training | Full training should be provided along with access to technical support |  | D
| Warranties | Warranty period should cover at least 12 months from installation |  | M
| Extended warranty | Extended warranty of up to 5 years |  | D
| Service support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. |  | D
| Service support | Details of preventative maintenance programmes required (if any) above Warranty requirements |  | D
OFFICE USE ONLY - This section below this line is for CPI Evaluation and Internal USE only
| Notes for Evaluation Team: |  | Name of Bidder to be inserted here | 1187 | Framework for the Supply of Chromatography Systems
Please assess all of the answers given, with particular attention to the Mandatory answers provided, taking into account, the Desirable and Optional answers too, however, you do not need to score each individual answer, but based on the scoring range 0-5, just give one singluar score (default set at "5" in the red and yellow box below, and as we have to provide feedback to bidders, you must also provide comments on which section(s) above where you felt that they did not meet the requirements.
| Evaluation Panel |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com