## Introduction
In addition to our original RFI submission please find further detailed requirements within this document The Full requirements for Each chamber can be found within the tabs provided Tab 1 CE 800-G16 Anechoic Chamber Tab 2 CE 800-106 Anechoic Chamber Tab 3 CE 800-106 Screen Rooms Tab 4 RF Tube B800 We would also like to request the answer to an additional question to our original RFI submission Question What are the implications, cost or otherwise, of installing the product in interim accommodation and moving it up to 12 months later to the final location? The interim and permanent locations are in buildings on the same site at Porton Down?

## T1-CE 800-G16 Anechoic Chamber
Helios Anechoic Chambers Room for Core Enclosure (CE) - Room 800-G16 Large Anechoic Chamber Room covering frequency band 26 MHz to 40 GHz V1.1B
Maximum External Dimensions | 10m x 6.8m x 5.7m (L x W x H)
Frequency Range | 26MHz to 40GHz
Measuring Distance | >3m
Maximum External Dimension of Ante-room | 2.6m x 2m x 2.2m (L x W x H)
Anechoic Chamber Usage Requirements
Radiate emission measurement
DEF STAN 59-411 Issue 3 | 10 kHz to 18GHz
EN55022 and CISPR CLASS B | 30 MHz to 1 GHz
CISPR 16-1-4 | 1 GHz to 18GHz
Immunity testing
Fully compliance IEC/EN 61000-4-3 | 80MHz to 6/18GHz
Other measurements
Antenna measurements | 30MHz to 40GHz
Radar RCS type measurements | >1 to 40GHz
Radio compatibility and EMC mitigation/investigations with RF communications | 30MHz to 40GHz
Chamber Performance Criterion
Maximum deviation from NSA to CISPR 16-1-4 | <± 3.5 dB
Maximum site VSWR to CISPR 16-1-4 | <± 6 dB
size of test volume | >=1.5m diameter/1.5 height
Size of Uniform area | 1.5m x 1.5m
Maximum deviation | 0dB/+6dB for 75% of 16 measuring points
RF Absorbers Hybrid
walls and ceiling | Full lining with Hybrid absorbers (non-critical, specular regions. Reflective edges to be covered)
Floor | Full lining with absorbers allowing emission and immunity, antenna and RCS related testing and facilitated walk absorbers to/from mast and turntable respectively.
Broadband performance | 26 MHz to 40 GHz with uniformity and isotropy . Detailed specification is to be provided with the proposal. Dstl may ask for the a copy of factory acceptance report and the performance will be a key factor in the decision to accommodate all varied users.
Product life time | >25 years, low maintenance , no carbon shredding over time, and sturdy or hard wearing and easily replaceable.
Power handling | >1 kW/m2 with the maximum field strength of >500 V/m
Fire-retardancy | NRL 8093 (tests 1,2 and 3), DIN4102 Class A or Class B, UL94HBF, ISO 11925-2 EURO Class E. Note that a test report and certificate are required for in-house Fire Safety officer at PTN.
White caps | For walls and ceiling
Clean room Standard | ISO 14644 Class 4 (non carbon shredding)
UV Certification | ISO 4892-2-2013
RoHs Compliant | 2011/65/EU
REACH Compliant | 1907/2006
Environmental | ISO/EN60068-2-1 and 60068-2-78.
Quality Control | ISO 9001/IEEE Standard 1128
Raised/Elevated Floor | Metallic ground plane height adjusted to the door step, with the maximum load capacity >1000 kg/m2, electrically conducted to the walls and ceiling panels. Ducting to accommodate all cabling from/to each wall patch panel to the specified floor patch panel. Also accommodate a movement of a tracked robot weighing up to 200kg.
Shielding Effectiveness 10kHz to 40GHz
Magnetic 10KHz | >55dB
Magnetic 100kHz | >75dB
Magnetic 1MHz-30MHz | >100dB
Plane wave 100MHz | >110dB
Plane wave 1 GHz | >110dB
Microwave 18 GHz | >75dB
Microwave 40 GHz | >75dB
Door
Door Size | >=1.5m x 2m (W XH) with a manual latching/unlatching with interlock switch and secure external locking facility.
Door Type | Single/multi leaf highly RF shielded door in TEMPEST and EMPP Environment.
Ventilation of the Chamber
Attenvents or honeycomb inserts to 40 GHz | Ventilation fans and control to keep the inside of the chamber at 20°C ±2 at the maximum radiated power <=6kW.
Option | Hot and cold air-conditioning system supplied and fitted
Electrical Installations >=100dB attenuation on load from 10 kHz to 40 GHz to MPE "Extended Performance Range" specification
Electrical system to comply with BS7671 safety standard
1 x Mains switch
3 x Emergency off switches (inside, outside and in Ante-room)
1 x Emergency light over the door
1 x Chamber use or indicator light when in use
The Lighting should have ultra-low noise emissions with 6 dB below DEF STAN 59-411 DRE01 limit (Plot is given below) requirement. A minimum illumination of 50 Lux must be achieved across the entire floor area. Dstl will request for a sample to be evaluated prior to the installation for its suitability.
Currently available 32A SPN
Currently available 32A TPN (Must supply two sockets in the chamber and one in Ante-room via raised floor). However, only one will be used at a time. Each socket will be in the corner at turntable and mast ends respectively.
If above stated mains supplies are not suitable then please specify whether you require instead 64A TPN and/or 32A TPN
DC 100V 60A for both chamber and Ante-room
Automated mast and turntable with controller
Turntable | 2m diameter, metallic surface electrically conducted to the ground-plane, height adjusted to the elevated ground-plane with the maximum load >=1000kg.
Mast | Automatic boresight mast, height adjustable to 1 to 4m including vertical and horizontal polarisation with an ability disable boresight option when required. Maximum weight 10-12kg. If not boresight mast them the maximum weight should be 15kg
Controller | Fibre optically controlling of mast and turntable facilitating the use of GPIB and USB
Device Under Test (DUT) and Safety Monitoring
Mobile wall mounted Video and voice camera | HD camera system with a pan/tilt controller facilitated with the recording option. DUT can be monitored by controlling the camera from ante-room or lab outside. Non invasive camera brackets/shelves need to be provided on walls. There will be total of 6 brackets and location TBD.
Small HD camera - Option | Small HD camera with a tripod to monitor close in DUT during test
RF Field probe(s) | Compact Tri-axial optical powered E field probe for monitoring RF field inside the anechoic chamber covering frequency band 10kHz to 8GHz.
Safety EMF monitoring/recording system providing continuous real-time measurements, which activates an audio/video alarm when EMF safety thresholds (CEMFAW16) are exceeded | 300kHz to 18GHz
Mobile wall mounted Infrared Red camera | Fixed Infrared camera opto-LWIR or similar specification with a pan tilt controller facilitated with the recording option. Likewise above, the same brackets/shelves will be used on walls.
All cameras should be EMC harden and have a high RF immunity of >200V/m to up to 40 GHz
Penetrations Panels(Exact location TBD) - All connection should be precision, female type and supplied with appropriate inside and outside tethered caps/covers
Inside the chamber to ante-room | 4 x N-Type, 4 x K-Type, 2 x SMA-Type, 2 x BNC-Type, 2 x 7/16-Type, 2 x LAN, 1 x Optical fibre feed through 8/6 ways, 1 x USB, 1 x penetration pipe (50mm diameter 300mm long with screw on cover) and RFI Trap.
Lab outside right-hand side of the doors | 4 x N-Type, 4 x K-Type, 2 x SMA-Type, 2 x BNC-Types, 2 x LAN, 1 x Optical fibre feed through 8/6 ways, 1 x USB, 1 x penetration pipe (50mm diameter 300mm long with screw on cover) and RFI Trap.
In centre Lab outside right hand side of the long wall | 2 x N-Type, 2 x K-Type, 2 x SMA, 2 x BNC-Types, 1 x LAN, 1 x Optical fibre feed through 6 ways, 1 x penetration pipe (50mm diameter 300mm long with screw on cover)
Floor Panels (Exact location TBD) - All connection should be precision female type and supplied with appropriate inside and outside tethered caps/covers
Behind Mast | 3 x N-Type, 3 x K-Type, 2 x SMA-Type, 1 x BNC-Type, 1 x 7/16-Type, 1 x Optical fibre feed through 4 ways and 2 x LAN
Behind Turntable and/or under turntable | 3 x N-Type, 3 x K-Type, 2 x SMA-Type, 1 x BNC-Types, 1 x 7/16-Type, 1 x Optical fibre feed through 4 ways and 2 x LAN
Turntable end in right hand corner | 2 x N-Type, 2 x K-Type, 1 x SMA-Type, 1 x BNC-Type and 1 x LAN
Mast end in left hand corner | 2 x N-Type, 2 x K-Type, 1 x SMA-Type, 1 x BNC-Type and 1 x LAN
Floor panels - Mains SPN, TPN and DC connection flush mounted sockets
Behind mast | 2 x Single phase Duplex 13A, 1 x single phase duplex non-RCD 13A, 1 x single phase 16A and 2 x DC 100V 60A
Behind Turntable (and/or under turntable) | 2 x Single phase Duplex 13A, 1 x single phase duplex non-RCD 13A, 1 x single phase 16A and 2 x DC 100V 60A
Turntable end on right hand corner | 1 x Single phase Duplex 13A and 1 x three phase 32A
Mast end on left hand corner | 1 x Single phase Duplex 13A and 1 x three phase 32A
On the centre of two long walls | 1 x Single phase Duplex 13A.
Ante-room | 1 x Three phase 32A (For RF amplifiers)
Turntable Ro-Jo/Slip rings (RF, mains and DC)
Duplex single phase 13A | From Ante-room and Lab outside right hand side of the door to Turntable
Single phase 16A | "
DC 100V 60A | "
N-type | "
K-Type | "
7/16-Type | From Ante-room to Turntable
Signal Connections (RF and DC) with an Ultra Low loss RF Cabling
From Ante- room to Turntable | N-Type, K-type, 7/16-Type, DC 100V 60A and LAN
From Ante-room to Behind mast | N-Type, K-type, 7/16-Type, DC 100V 60A and LAN
From lab outside right hand side of the door to Turntable | N-Type, K-type, DC 100V 60A and LAN
From lab outside right hand side of the door to Behind mast | N-Type, K-type, DC 100V 60A and LAN
Broadband (swept) Anechoic Chamber Calibration and Validation
Shielding effective 10kHz to 40 GHz prior and post installation of the ferrite tile and absorbers | MIL STD 285 or EN 50147-1
Normalised Site Attenuation (NSA) | EN50147-2 or CISPR 16-1-4
Site VSWR measurements | EN50147-2 or CISPR 16-1-4
Filed uniformity | EN 61000-4-3 to 6GHz or with an option to test to 18GHz
Ante-Room Specification
Maximum external dimension each Screen Room | 2.6m x 2m x 2.2m (Lx W x H)
Frequency Range | 10kHz to 40GHz
Shielding Effectiveness 10kHz to 40 GHz
Magnetic 10KHz | >55dB
Magnetic 100kHz | >75dB
Magnetic 1MHz-30MHz | >100dB
Plane wave 100MHz | >100dB
Plane wave 1 GHz | >100dB
Microwave 18 GHz | >75dB
Microwave 40 GHz | >75dB
Doors
Door Size | 1.2 m x 2m (W XH) with a manual latching/unlatching
Door Type | Single/multi leaf highly RF shielded
Ventilation of each screen room
Attenvents or honeycomb inserts to 40 GHz | Ventilation fans and control to keep the inside of the chamber at 20°C ±2 at the maximum radiated power <=6kW.
Electrical Installations >=100dB attenuation on load from 10 kHz to 40 GHz to MPE "Extended Range" specification
Electrical system to comply with BS7671 safety standard
1 x LED lighting
1 x Mains switch
1 x Emergency cut-off switch inside
1 x Emergency exit light over the door
1 x chamber room use or indicator light outside when in use
Currently available two of 32A SPNs
DC 100V 60A
If above stated mains supplies are not suitable then please specify whether you require instead 64A TPN and/or 32A TPN
All connector in penetration panels should be precision, female, and with inside and outside tethered caps/covers - Exact location TBD
Inside ante-room to the chamber | 4 x N-Type, 4 x K-Type, 2 x SMA-Type, 2 x BNC-Type, 2 x 7/16-Type, 2 x LAN, 1 x Optical fibre feed through 8/6 ways, 1 x USB, 1 x penetration pipe (50mm diameter 300mm long with screw on cover) and RFI Trap.
Right hand site side of the wall | 4 x N-precision, 4 x K-Type, 3 x SMA, 2 x BNC-Type, 2 x LAN, 1 x Optical fibre feed through 8/6 ways, 1 x USB, and 1 x penetration pipe (50mm diameter 300mm long with screw on cover)
Miscellaneous filter specification for Ante-room and chamber
32A SPN and TPN | >=100dB attenuation on load from 10kHz to 40GHz to MPE "Extended Range" specification
6A No Door Entry Supply | "
LAN Filters | "
Alarm light Filter | "
Emergency Voice alert Filter | "
Phone Filter | "
Broadband Shielding effective testing - Validation
Shielding Effectiveness 10kHz to 40 GHz | Broadband (sweep) either MIL-STD 285 or EN 50174-1
Miscellaneous Items
Ante-room | Raised floor in line with the door lips with lining and outside ramps to be provided
Styrofoam Table | DUT 1.5m diameter round (or square), low permitivity, table
GPS | Repeater (GSP signal is available on lab benches)
Option | Set of antennas covering band 30MHz to 40GHz
Option | RF amplifier(s) to 30MHz to 3GHz
Compressed air | 1 x compressed air feed through if a selected mast requires this
Post Helios Completion Tasks by FMS/EMCOR (TBD)
To Provide | 32ATPN supply and alternative to chamber 64A TPN/32A TPN if requested by a supplier
Installation of mains sockets and benches inside Ante-room | Number of sockets on each long walls xx and Number of sockets on each short walls xx
Phone | 1 x Ante-room
Emergency voice alert | 1 x Ante-room and the chamber
Connection of 32A SPN and TPN mains | To Ante-room and anechoic chamber
If hot and cold air conditioning option is selected | Appropriate ventilation fitting to chamber vents
Fire detection sensors | Anechoic chamber and ante-room attenvents
Lighting emissions must be at least 6dB below DEF STAN 59-411 DRE01 Limits
CE 800 -G16 Anechoic Camber and Ante-room details with potential location of each wall and floor penetration panels (TBD)

## T2-CE 800-106 Anechoic Chamber
Helios Anechoic Chamber Room for Core Enclosure (CE) - Room 800-106 High Power Anechoic Chamber Room Covering frequency band 26/80MHz to 40GHz V1.1B
Maximum External Dimension | 7.0m x 6.2m x 3.0m (Lx W x H)
Frequency Range Option 1 | 30MHz to 40GHz Hybrid
Frequency Range Option 2 | 80MHz to 40GHz Standard and only ceiling hybrid with the absorber length <= 450mm and other absorbers <= 660mm length.
Measuring Distance | 3.0m
Maximum External Dimension of Ante-room | 3.7m x 3.125m x 2.2m
Anechoic Chamber Usage Requirements
Radiate emission measurement
DEF STAN 59-411 Issue 3 | 10kHz to 18GHz
EN55022 and CISPR CLASS B | 30MHz or 80MHz to 1GHz
CISPR 16-1-4 | 1GHz to 18GHz
Immunity testing
Fully compliance IEC/EN 61000-4-3 | 80MHz to 6/18GHz
Other measurements
Antenna measurements | 30MHz or 80MHz to 40GHz
Radar RCS type measurements | >1GHz to 40GHz
Radio compatibility and EMC mitigation/investigations with RF communications | 30MHz to 40GHz
Chamber Performance Criterion
Maximum deviation from NSA to CISPR 16-1-4 | <=± 3.5dB
Maximum site VSWR to CISPR 16-1-4 | <=± 6dB
size of test volume | >=1.0 m diameter/>=1.0m height
Size of Uniform area | >= 1.0 m x 1.0 m
Maximum deviation | 0dB/+6dB for 75% of 16 measuring points
RF Absorbers Hybrid or Standard
Walls and ceiling | Full lining with Hybrid/standard absorbers (non-critical, specular regions and reflective edges to be covered)
Floor | Full lining with absorbers allowing emission and immunity, antenna and RCS related testing and facilitated walk absorbers to/from mast and turntable respectively.
Broadband performance | 26MHz/80MHz to 40GHz with uniformity and isotropy. Detailed specification is to be provided with the proposal. Dstl may ask for a copy of factory acceptance report and the performance will be a key factor in the decision to meet the requirements of all stated users.
Product life time | >25 years, low maintenance , no carbon shredding over time, and sturdy or hard wearing and easily replaceable.
Power handling | >=1 kW/m2Hybrid or 1.5 kW/m2with the maximum field strength of >500V/m
Fire-retardancy | NRL 8093 (tests 1,2 and 3), DIN4102 Class A or Class B, UL94HBF, ISO 11925-2, EURO Class E. Note that a test report and certificate are required for in-house Fire Safety officer at PTN site.
White caps | For walls and ceiling
Clean room Standard | ISO 14644 Class 4 (non carbon shredding)
UV Certification | ISO 4892-2-2013
RoHs Compliant | 2011/65/EU
REACH Compliant | 1907/2006
Environmental | ISO/EN60068-2-1 and 60068-2-78
Quality Control | ISO 9001/IEEE Standard 1128
Raised/Elevated Floor | Metallic ground plane height adjusted to the door step, with the maximum load capacity >=1000 kg/m2, electrically conducted to the walls and ceiling panels. Ducting under the Ground-plane enabling cabling run to/from each specified wall patch panel to each floor patch panel. Also accommodate a movement of a tracked robot weighing up to 200kg.
Shielding Effectiveness 10kHz to 40GHz
Magnetic 10KHz | >55dB
Magnetic 100kHz | >75dB
Magnetic 1MHz-30MHz | >100dB
Plane wave 100MHz | >110dB
Plane wave 1GHz | >110dB
Microwave 18GHz | >75dB
Microwave 40GHz | >75dB
Door
Door Size | >=1.5m x 2m (W x H) with a manual latching/unlatching with interlock switch and secure external locking facility.
Door Type | Single/multi leaf highly RF shielded door environment.
Ventilation of the Chamber
Attenvents or honeycomb inserts to 40 GHz | Ventilation fans and control to keep the inside of the chamber at 20°C ±2 at the maximum radiated power <=6kW.
Option | Hot and cold air-conditioning system supplied and fitted
Electrical Installations >=100dB attenuation on load from 10 kHz to 40 GHz to MPE "Extended Performance Range" specification
Electrical system to comply with BS7671 safety standard
1 x Mains switch
3 x Emergency off switches (inside, outside and in Ante-room)
1x Emergency light over the door
1 x Chamber use or indicator light when in use
The Lighting should have ultra-low noise emissions with 6 dB below DEF STAN 59-411 DRE01 limit (Plot is given below) requirement. A minimum illumination of 50 Lux must be achieved across the entire floor area. Dstl will request for a sample to be evaluated prior to the installation for its suitability.
Currently available 64A SPN
Currently available 32A TPN
Currently available 32A TPN (Must supply two sockets in the chamber and one in Ante-room via raised floor). However, only one will be used at a time. Each socket will be in the corner at turntable and mast ends respectively.
If above stated mains supplies are not suitable then please specify whether you require instead 64A TPN and/or 32A TPN
DC 100V 60A for both chamber and Ante-room
Automated Mast, Turntable and Controller
Turntable | Innco or Innco compatible controller 1.5m diameter, flush mounted, low noise, metallic surface electrically conducted to the ground-plane, height adjusted to the elevated ground-plane with the maximum load >=1000kg.
Mast option | Innco or Innco compatible controller, automatic boresight mast, height adjustable to 1 to 2m height and including vertical and horizontal polarisation with an ability to disable boresight option as when required. The maximum weight 10-12kg.
Controller option | Fibre optically controlling of the mast and turntable with GPIB and USB control.
Device Under Test (DUT) and Safety Monitoring
Mobile wall mounted HD video and voice camera | mk-messtechnik dAV-Cr-HD-30-nd, HiRF HD camera system with pan/tilt controller and facilitated with the recording option. DUT can be monitored by controlling the camera from ante-room or lab outside. Non invasive camera brackets/shelves need to be provided on walls. There will be total of 6 brackets and location TBD.
Mobile wall mounted Infrared Red camera | mk-messtechnik Infrared camera opto-LWIR with a pan tilt controller facilitated with the recording option. Likewise above, the same brackets/shelves will be used on walls.
RF Field probe(s) | Compact Tri-axial optical powered E field probe for monitoring RF field inside the anechoic chamber covering frequency band 10kHz to 8 GHz.
Safety EMF monitoring/recording system providing continuous real-time measurements, which activates an audio/video alarm when EMF safety thresholds (CEMFAW16) are exceeded | 300 kHz to 18GHz
RF Immunity | All cameras should be EMC hardened and have a high RF immunity of >600V/m up to 40 GHz, 40kV/m in small band at 3 GHz and >90kV/m UWB pulse.
Penetrations Panels(Exact location TBD) - All connection should be precision, female type and supplied with appropriate inside and outside tethered caps/covers
Inside the chamber to ante-room | 4 x N-Type, 4 x K-Type, 2 x SMA-Type, 2 x BNC-Type, 2 x 7/16-Type, 2 x LAN, 1 x Optical fibre feed through 8/6 ways, 1 x USB, 1 x penetration pipe (50mm diameter 300mm long with screw on cover) and RFI Trap.
Lab outside right-hand side of the doors | 4 x N-Type, 4 x K-Type, 2 x SMA-Type, 2 x BNC-Type, 2 x LAN, 1 x Optical fibre feed through 8/6 ways, 1 x USB, 1 x penetration pipe (50mm diameter 300mm long with screw on cover) and RFI Trap.
In centre Lab outside right hand side of the long wall | 2 x N-Type, 2 x K-Type, 2 x SMA, 2 x BNC-Type, 1 x LANS, 1 x Optical fibre feed through 6 ways and 1 x penetration pipe (50mm diameter 300mm long with screw on cover)
Floor Panels (Exact location TBD) - All connection should be precision female type and supplied with appropriate inside and outside tethered caps/covers
Behind Mast | 3 x N-Type, 3 x K-Type, 2 x SMA-Type, 1 x BNC-Types, 1 x 7/16-Type, 1 x Optical fibre feed through 4 ways and 2 x LAN
Behind Turntable and/or under turntable | 3 x N-Type, 3 x K-Type, 2 x SMA-Type, 1 x BNC-Types, 1 x 7/16-Type, 1 x Optical fibre feed through 4 ways and 2 x LAN
Turntable end in right hand corner | 2 x N-Type, 2 x K-Type, 1 x SMA-Type and 1 x BNC-Type and 1 x LAN
Mast end in left hand corner | 2 x N-Type, 2 x K-Type, 1 x SMA-Type and 1 x BNC-Type and 1 x LAN
Floor panels - Mains SPN, TPN and DC connection flush mounted sockets
Behind mast | 2 x Single phase Duplex 13A, 1 x single phase duplex non-RCD 13A, 1 x single phase 16A and 2 x DC 100V 60A
Behind Turntable (and/or under turntable) | 2 x Single phase Duplex 13A, 1 x single phase duplex non-RCD 13A, 1 x single phase 16A and 2 x DC 100V 60A
Turntable end on right hand corner | 1 x Single phase Duplex 13A and 1 x three phase 32A
Mast end on left hand corner | 1 x Single phase Duplex 13A and 1 x three phase 32A
On the centre of two long walls | 1 x Single phase Duplex 13A.
Ante-room | 1 x Three phase 32A (For RF amplifiers)
Turntable Ro-Jo/Slip rings (RF, mains and DC)
Duplex single phase 13A | From Ante-room and Lab outside right hand side of the door to Turntable
Single phase 16A | "
DC 100V 60A | "
N-type | "
K-Type | "
7/16-Type | From Ante-room to Turntable
Signal Connections (RF and DC) with an Ultra Low loss RF Cabling
From Ante- room to Turntable | N-Type, K-type, 7/16-Type and DC 100V 60A and LAN
From Ante-room to Behind mast | N-Type, K-type, 7/16-Type and DC 100V 60A and LAN
From lab outside right hand side of the door to Turntable | N-Type, K-type and DC 100V 60A and LAN
From lab outside right hand side of the door to Behind mast | N-Type, K-type and DC 100V 60A and LAN
Broadband (swept) Anechoic Chamber Calibration and Validation
Shielding effective 10kHz to 40 GHz prior and post installation of the ferrite tile and absorbers. | MIL STD 285 or EN 50147-1
Normalised Site Attenuation (NSA) | EN50147-2 or CISPR 16-1-4
Site VSWR measurements | EN50147-2 or CISPR 16-1-4
Filed uniformity | EN 61000-4-3 to 6GHz or with an option to 18GHz
Ante-Room Specification CE 800-106
Maximum External Dimension | 3.7m x 3.125m x 2.2m (Lx W x H)
Frequency Range | 10kHz to 40 GHz
Shielding Effectiveness 10kHz to 40 GHz
Magnetic 10KHz | >55dB
Magnetic 100kHz | >75dB
Magnetic 1MHz-30MHz | >100dB
Plane wave 100MHz | >100dB
Plane wave 1GHz | >100dB
Microwave 18GHz | >75dB
Microwave 40GHz | >75dB
Door
Door Size | 1.5 m x 2m (W X H) with a manual latching/unlatching
Door Type | Single/multi-leaf highly RF shielded
Ventilation of the Ante-room
Attenvents or honeycomb inserts to 40 GHz | Ventilation fans and control to keep the inside of the chamber at 20°C ±2 at the maximum radiated power <=6kW.
Electrical Installations >=100dB attenuation on load from 10 kHz to 40 GHz to MPE "Extended Range" specification
Electrical system to comply with BS7671 safety standard
1 x LED lighting
1 x Mains switch
1 x Emergency cut-off switch inside
1 x Emergency exit light over the door
1 x Ante-room use or indicator light outside when in use
Currently available 2 off 32A SPNs
DC 100V 60A
If above stated mains supplies are not suitable then please specify whether you require instead 64A TPN and/or 32A TPN
All connector in penetration panels should be precision, female, and with inside and outside tethered caps/covers - Exact location TBD
Inside ante-room to the chamber | 4 x N-Type, 4 x K-Type, 2 x SMA-Type, 2 x BNC-Type, 2 x 7/16-Type, 2 x LAN, 1 x Optical fibre feed through 8/6 ways, 1 x USB, 1 x penetration pipe (50mm diameter 300mm long with screw on cover) and RFI Trap.
Right hand site side of the wall | 4 x N-precision, 4 x K-Type, 3 x SMA, 2 x BNC-Type, 2 x LAN, 1 x Optical fibre feed through 8/6 ways, 1 x USB, and 1 x penetration pipe (50mm diameter 300mm long with screw on cover).
Miscellaneous filter specification for Ante-room and chamber
32A SPN and TPN | >=100dB attenuation on load from 10kHz to 40GHz to MPE "Extended Range" specification
6A No Door Entry Supply | "
LAN Filters | "
Alarm light Filter | "
Emergency Voice alert Filter | "
Phone Filter | "
Broadband Shielding effective testing - Acceptance
Shielding Effectiveness 10kHz to 40 GHz | Broadband (swept) either MIL-STD 285 or EN 50174-1
Miscellaneous items
Ante-room | Raised floor in line with the door lips with lining and outside ramps to be provided
Styrofoam Table | DUT 1.0m diameter round (or square), low permittivity, table
GPS | Repeater (GSP signal is available on lab benches)
Option | Antenna measurement software
Compressed air | 1 x compressed air feed through if a selected mast requires this
Post Helios Completion Tasks by FMS/EMCOR (TBD)
To Provide | 32ATPN supply and alternative to chamber 64A TPN/32A TPN if requested by a supplier
Installation of mains sockets and benches inside Ante-room | Number of sockets on each long walls xx and Number of sockets on each short walls xx
Phone | 1 x Ante-room
Emergency voice alert | 1 x Ante-room and the chamber
Connection of 32A SPN and TPN mains | To Ante-room and anechoic chamber
If hot and cold air conditioning option is selected | Appropriate ventilation fitting to chamber vents
Fire detection sensors | Anechoic chamber and ante-room attenvents
Lighting emissions must be at least 6dB below DEF STAN 59-411 DRE01 Limits
CE 800 -106 Anechoic Camber and Ante-room details with potential location of each wall and floor penetration panels (TBD)

## T3-CE 800-106 Screen Rooms
Helios Screen Rooms for Core Enclosure (CE) - Room 800-106 - Two Back-to-Back Connected Screen Rooms
Maximum external dimension each Screen Room | 3m x 3.5m x 3m (W x L x H)
Frequency Range | 10kHz to 40GHz
Shielding Effectiveness 10kHz to 40GHz
Magnetic 10KHz | >55dB
Magnetic 100kHz | >75dB
Magnetic 1MHz-30MHz | >100dB
Plane wave 100MHz | >100dB
Plane wave 1GHz | >100dB
Microwave 18GHz | >75dB
Microwave 40GHz | >75dB
Doors (Total of 3 doors: Two external and one internal between the screen rooms)
Each Door Size | 1m x 2m (W X H) with a manual latching/unlatching with interlock switch for the external doors.
Door Type | Single/multi leaf highly RF shielded
Ventilation of each screen room
Attenvents or honeycomb inserts to 40GHz | Ventilation fans and control to keep the inside of the chamber at 20°C ±2 at the maximum radiated power <=2kW.
Electrical Installations
For each screen room | Electrical system to comply with BS7671 safety standard
1 x Low noise version of LED lighting
1 x Mains switch
2 x Emergency off switches (one inside and one outside the screen room)
1 x Emergency exit light over the door
1 x screen room use or indicator light outside when in use
Available 32A SPN
Penetration Panels (Exact Location TBD)- All connections should be precision female and with appropriate inside and outside tethered caps/covers - Probable locations specified in the attached drawing
Each patch panel outside | 4 x N, 4 x K-Type, 4 x SMA-Type, 4 x BNC-Type, 2 x LAN, 1 x Optical fibre feed through 6 ways, 1 x penetration pipe (50mm diameter 300mm long with screw on cover) and RFI trap.
Patch panel between two screen rooms | 4 x N, 4 x K-Type, 4 x SMA-Type, 4 x BNC-Type, 2 x LANS and 1 x Optical fibre feed through 6 ways.
Filter Specification for Each Screen Room
Single phase 32A | >=100dB attenuation on load from 10kHz to 40GHz to MPE "Extended Range" specification
6A No Door Entry Supply | "
LAN Filters | "
Alarm light Filter | "
Emergency Voice alert Filter | "
Phone Filter | "
Broadband Shielding Effective Testing - Validation
Shielding Effectiveness 10kHz to 40GHz | Broadband (swept) either MIL-STD 285 or EN 50174-1
Miscellaneous
Each screen room | Raised floor in line with the door lips and outside ramps to be provided and flat shield and floor lined.
Post Helios Completion of Tasks by FMS/EMCOR TBD
Installation mains sockets and benches inside each screen room | Number of sockets on each long walls xx and Number of sockets on each short walls xx
Phone | 1 x in a screen room
Emergency voice alert | 1 x in each screen room
Fire detection sensors | Screen room attenvents
Available 32A SPN mains | connection to each screen room

## T4-RF Tube B800
Design Specification Anechoic chamber requirements: • Operating frequency range: 500 MHz – 12 GHz • Shielding isolation: <100 dB • External chamber dimensions: Width 1500, Depth 4000, Height 2000 (mm) • Doors approx. 650 x 650 (mm) are required at each end of the chamber • Penetration panels at each end of the chamber, each with 4 x SMA-Type and 4 x N-Type connectors (All connections should be precision, female type and supplied with appropriate inside and outside tethered caps/covers ) • Internal fibre lighting is required (No LEDs) • All chamber parts must be able to fit through standard double doors as the chamber will be constructed on site (ground floor) • VHP-12 RAM standard/hybrid (chosen for optimum performance over operating frequency range), non-carbon shredding, fire retardant • Fire retardant standard: NRL 8093 (tests 1,2 and 3), DIN4102 Class A or Class B, UL94HBF, ISO 11925-2, EURO Class E. Note that a test report and certificate are required for in-house Fire Safety officer at PTN site. • Bid must include a wheeled stand suitable to support the chamber and raise it to a working height • Mains filtering >=100dB attenuation from 10 KHz to 40 GHz. Remote turntable/platform requirements: • Turntable must be able to move backwards and forwards down the useable length of the chamber in 1 mm steps. Must be able to traverse the length of the chamber in approx. 2 minutes • Turntable must be able to rotate 360 degrees in 1 degree steps (full rotation in less than a minute) • All motors, electronics, and metal need to be external to the chamber. Only plastic, nylon, rubber, etc can be inside the chamber. • It is NOT acceptable to have shielded electronics inside the chamber. • For use with payloads up to 2.5 kg • The design should not be belt driven. • Must be remote operated by serial commands or similar