HSE/T3618

STATEMENT OF SERVICE REQUIREMENTS FOR THERMOCOUPLE CALIBRATION UNITS : REQUEST FOR QUOTE

- HEALTH & SAFETY EXECUTIVE, SCIENCE DIVISION

1.1 	The Health and Safety Executive (HSE) is a Crown non-departmental public body with specific statutory functions in relation to health and safety.  It is appointed by the Secretary of State for Work and Pensions and employs around 3500 staff including policy advisers, inspectors, technologists and scientific and medical advisers. HSE’s job is to prevent people being killed, injured or made ill by work.

  - HSE consists of a governing Board comprising of a Chair and nine non-executive members, all of whom are appointed by the Secretary of State for Work and Pensions after consultation with organisations representing employers, employees, local authorities and others, as appropriate.  HSE comprises various directorates and groups and is led by a senior management team.  HSE`s Science & Research Centre is located within a 550 acres site near Buxton in Derbyshire and is part of HSE`s Science Division.  HSE works from over 30 locations throughout Great Britain.

  - The science and research centre is Britain's leading industrial & occupational health and safety research facility. Science Division supports HSE’s mission to protect the health and safety of the national workforce by ensuring risks in the workplace are adequately controlled.

  - Science Division employs circa 420 people including scientists, engineers, psychologists, social scientists, health professionals and technical specialists.  Its capabilities encompass a wide range of topics including: health solutions, risk and human factors, fire, explosion and process safety; occupational and environmental health; safety engineering; work environment; and specialist photographic and technical services.

Services Include:

- Research and development

- Specialist advice and consultancy

- Forensic investigation into the causes of accidents

- Environmental and biological monitoring

- Assessment of levels of risk and investigation of their control

- Establishing realistic requirements for standards, and processes or meeting those standards

- Validation and certification

- Training

  - To deliver these services Science Division has advanced facilities that range from high power computers for modeling and analysis, well equipped laboratories covering biomedical, occupational hygiene and environmental work to unique facilities for large scale experiments in the areas of engineering, fires, explosions and process safety.

- BACKGROUND TO THE PROJECT

Temperature calibrator specification

Science Directorate requires a temperature calibration system for verifying and calibrating multiple thermocouples and PRTs across a wide temperature range to a high degree of accuracy. The specification below outlines our base requirements.

A single or combination of mobile dry block calibrators covering a temperature range from -50 °C to 1200 °C. These calibrators should:

•	Be mobile/portable with independent indicators and carry handle/s

•	Be 110 Vac compatible

•	Be supplied with compatible reference TCs/PRTs

•	Be supplied with protective carry case/s

•	Be calibrated to UKAS standards (minimum 5 point calibration)

•	Have suitable dry blocks to accommodate up to 8 sensors plus a reference for the following sensor diameters: 0.5 mm, 1 mm, 1.5 mm, 2 mm, 3 mm

•	Have blank inserts

•	Have a minimum full range temperature stability of 0.2 °C

•	Have a minimum full range temperature resolution of 0.1 °C

•	Have a heating rate ≥25 °C per minute

•	Have a cooling rate from 1200 to 200 °C of ≤60 minutes (cooling adapter would be beneficial)

•	Ceramic insulators for top of each dry block

A precision thermometer unit capable of interfacing up to 8 sensors (including PRTs, thermocouples and 4-20 mA inputs) simultaneously. This precision thermometer should have:

•	A resolution of 0.0001 °C

•	An initial accuracy of 5 mK for PRTs, 2 µV for TCs and 0.002 mA for 4-20 mA outputs

•	The ability to control the dry block calibrator/s through a series of temperatures or cycles

•	Have internal or external USB logging

•	Software compatibility with the temperature calibration package

•	Cables to connect to windows PC

•	UKAS calibration to the best uncertainty required (minimum 5 point calibration)

•	A protective carry case/s

Calibration software package with one license seat. Software should have the ability to perform automatic sensor calibration for up to 16 sensors and provide uncertainty coefficients and customised calibration certificates. The results of the calibrations should be saveable to file.

We have two ISOTECH dry block calibrators which could be traded in for a 20% discount with the manufacturer should that be required. These are currently in need of repair and will be superseded by the new unit/s.

- Request for Quote Response

Please provide a quote for the goods requested and  confirm that you can meet the delivery date of Friday 15th May 2020.