## Sheet1
USER REQUIREMENT for: |  | Tribological Testing Capability |  |  | Tenderer's Response
Requirement codes
Q | Mandatory -linked to the Critical Quality Attributes & Critical Process Parameters of the process and will be tested during the qualification of the system
S | Mandatory Safety Requirement (Bio, toxicity, flammable, manual handling etc)
M | Mandatory, but not a GMP or Qualification test
D | Desirable, but not necessary - if this cannot be supplied. Tenderer should suggest alternative if appropriate
N | Not required or not applicable
| Parameter | Answer | Requirement (Q/S/M/D/N) -see codes | Notes | Answer or Details | Conforms? (Y/N)
Scope
Range of use |  | A single system which has the ability to measure friction, wear, and scratch, in a range of contact configurations under a range of environemental testing conditions. | Q
|  | A single system which can be configured to perform ASTM, ISO and DIN Tribology and mechanical testing methods (such as D6079, D4518, G133, G099, D2266, G2981, G171, E18) | Q
|  | Must include a linear drive, a fast reciprocating linear drive, and a high speed rotarty drive in the single system. | Q
|  | In addition to those features suggested above the system must measure multi-axi strain properties including tension and compression. | Q
|  | The system must test under recipocating motion with a stroke range of at least 0.1 mm to 25 mm, with test speeds up to 60 Hz, with test loads up to 2 kN. | Q
|  | The system must have variable rolling-to-sliding contact for rotary motion. | Q
|  | Interchangeable and quick connect 2 axis force sensors ranging from 1mN to 2kN. The system must include 1 to 100 mN ; 0.05 to 5 N and 1 to 100 N force sensors. | Q
Environmental Capability
|  | It is essential that samples can be tested in both a dry and liquid environment for reciprocating motion from ambient up to 400 deg. | Q
|  | It is essential the samples can be tested in both a dry and liquid environment for rotary motion from ambient up to 400 deg. | Q
|  | It is essential that samples can be tested at below ambient temperature ( down to -25deg) for rotary motion. | Q
General |  | Versatility
|  | Design which enables additional functionalities to be added | D
|  | Interchangeable and tool-less changeover of contact configurations. | D
|  | Additional BNC connections for custom analog sensor integration | D
Ease of use
|  | The system should be supplied complete, including computer with unlimited licences to put the software on a number of computers. | Q
|  | The software should be easy to operate. The data processing should be straight forward to operate by someone familiar with tribological/mechanical testing. | Q
|  | Automatic load protection should be included | Q
|  | Should include built-in software for force sensor calibration | Q
Fixtures
|  | The system needs to be supplied with a range of component holders for trioboligcal testing including pin/ball on disc (range of sized pins/balls), disc/plate-on-disc/plate, indenter on plate, crossed cylinders, 4 ball test, and diamond stylus tips. This must include a range of fixtures to allow pin/ball on plate/disc for use with high temperature (up to 400deg), and low temperature (down to -25deg) chambers, for the range of load cells provided. | D
Dimensions |  | Compact bench top laboratory design | Q
Performance
|  | Built in temperature controller | Q
|  | Closed loop servo-controlled z-axis | Q
|  | Auto-recognition system to automatically link which modular system is being used | D
|  | Pre-loaded frequently used software controlled tribological standards which can be optimized for specific testing | D
Options |  | Acoustic-emissions and recording sensors | D
|  | Electrical contact resistance | D
|  | Programmable Pump for controlled fluid testing | D
|  | Optical microscopy with multiple objective lenses (x5,x10x20 objectives with 1280x1024 digital camera) | D
|  | A range of consumables such as pins/plates/balls/indenters in a range of sizes/dimensions. | D
|  | In situ wear depth monitoring | D
Guidance: Tenderers must provide responses within the following Columns;
1. Column G - Please provide details of your equipment functionality in relation to CPI requirements specified in column D
2. Column H - Please confirm if your product complies with CPI's functionality requirements specified in Column D, by providing a Yes (Y) or No (N) response