Schedule A

### THE HEALTH AND SAFETY EXECUTIVE

STATEMENT OF REQUIREMENTS FOR THE PROVISION OF SUPPLY OF A HYDROGEN GENERATOR T21 1516

1	THE HEALTH AND SAFETY EXECUTIVE (HSE)

1.1	The Health and Safety Executive (HSE) is a Crown non-departmental public body with specific statutory functions in relation to health and safety.  It is appointed by the Secretary of State for Work and Pensions and employs around 3500 staff including policy advisers, inspectors, technologists and scientific and medical advisers. HSE’s job is to prevent people being killed, injured or made ill by work.

  - HSE consists of a governing Board comprising of a Chair and nine non-executive members, all of whom are appointed by the Secretary of State for Work and Pensions after consultation with organisations representing employers, employees, local authorities and others, as appropriate.

1.3	HSE comprises various directorates and groups and is led by a senior management team.  The Health and Safety Laboratory (HSL), Chemicals Regulation Directorate (CRD) and Office for Nuclear Regulation (ONR) are also part of HSE.  HSE works from over 30 locations throughout Great Britain.

2.	HEALTH & SAFETY LABORATORY

The Health and Safety Laboratory (HSL) is Britain's leading industrial health and safety facility. Operating as an agency of the Health and Safety Executive (HSE), HSL supports HSE’s mission to protect people's health and safety by ensuring risks in the changing workplace are properly controlled.

The Health & Safety Laboratory are inviting tenders for the supply of a Hydrogen Generator suitable for gas chromatographic (GC) applications.

Hydrogen generators are used at HSL to eliminate hydrogen gas lines and cylinders and supply high purity hydrogen to GC equipment as both carrier gas and for flame ionization detectors (FIDs).  They are also used to supply hydrogen to standalone FIDs. Existing hydrogen generators are used to supply up to four GC instruments. For the purposes of this invitation to tender it is necessary for HSL to be open and fair to possible sources that believe they can meet these requirements and may submit a tender.

3	BACKGROUND TO THE PROJECT

3.1	Hydrogen generators are used at HSL to eliminate hydrogen gas lines and cylinders and supply high purity hydrogen to GC equipment as both carrier gas and for flame ionization detectors (FIDs).  They are also used to supply hydrogen to standalone FIDs. Existing hydrogen generators are used to supply up to four GC instruments. For the purposes of this invitation to tender it is necessary for HSL to be open and fair to possible sources that believe they can meet these requirements and may submit a tender.

4.	SPECIFICATION REQUIRED

4.1	The supply of the instrument will include all hardware components and appropriate electrical and gas connections.

Specification for Hydrogen Generator

System Features

Capable of delivering at least 250 ml/min of hydrogen at a pressure of up to 60 psi

Gas Purity > 99.999%

Automatic cut off if a hydrogen leak is detected

4.2	Installation, Commissioning, Training

Tender to include installation commissioning and delivery to HSL Buxton Derbyshire

and on-site training.

4.3	Warranty and Documentation

Customary warranty and / or extended warranty to be offered.

All user manuals should be received by HSL with the system hardware.

On warranty HSL would like to achieve 24 month warranty period for the instrument as standard.

Please provide warranty information, stating whether it is standard warranty (return to base) or onsite warranty (next business day), together with the length of warranty that will be supplied with the proposed equipment. Please also provide details of exactly how these warranties shall be activated and/or requested, should any problems occur.

5.	Information Technology

HSL has a separate Facilities Management IT Provider whom are responsible for the whole network of the laboratory.  In that regard, they have a clear responsibility for ‘impact testing’ new IT equipment onto HSL’s IT network. The laboratory’s FM IT Provider sets a minimum requirement on all new equipment.

It is therefore a requirement that all Tenderers should specify in the fullest detail possible, the makers name and specification etc of all IT equipment (hardware, software etc.) that they propose to supply including details of the proposed operating system. Details of any associated IT peripherals and software should also be supplied. It would greatly assist if separate itemised costs for all of the associated IT components (e.g. PC, LCD, and Printer) could be included within the tender response, should these be available.

In addition, please provide warranty information, stating whether it is standard warranty (return to base) or onsite warranty (next business day), together with the length of warranty that will be supplied with the proposed equipment. Please also provide details of exactly how these warranties shall be activated and/or requested, should any problems occur.

PC Minimum Specification:

Processor: Intel i5 (or equivalent)

Memory: 4GB

Hard disk: 500GB

Display: DVI capable of 1280 x 800 resolution

Operating System: Windows 7 Professional

Connectivity: Network adapter (Ethernet RJ-45)

The PC is considered as part of the 'system' we buy and should be specifically included in any warranty agreement offered by the tenderer. We would expect that any problems or repairs with the software and/or hardware that forms part of the 'system' should be resolved by the supplier.

6.	Software

Prior to placing any order, HSL requires assurance from the vendor that the proposed software contained within the order does not require administration rights to run and operate correctly post installation. This is because HSL is unable to run and operate software in MS Windows that needs to have administration rights before it will run and operate correctly post installation. Therefore we wish the vendor to:

- Confirm that the software does not require administration rights to run; or

- If you are unable to provide a version of software that does not run and operate correctly without “administration rights”, please advise what you can do to resolve this.

Note that in the event of an order being placed, and it being found that “administration rights” are required to run and operate the software, or, alternatively, that you do not notify HSL's Purchasing Procurement Unit that the software needs administration rights in order to run and operate correctly, HSL reserves the right to cancel the order and to place an order with another supplier and to seek a refund of any payments that have been made against the purchase order.