## Drop Codes (hidden)
RoomClassification | IA | QualificationTest
Not Classified | Direct Impact | Yes
Grade A | Indirect Impact | No
Grade B | No Impact | n/a
Grade C
Grade D

## Multimode_Plate_Reader
User Requirement Specification |  |  | CPI Internal Tender Reference & Title:- |  | Multimode plate reader (for research use only)
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 35%
USER REQUIREMENT |  |  |  |  | Tenderer Response
Requirement codes |  |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as Desirable) |  |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  | Requirement |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
M1 | Multimode plate reader | Able to process 6 to 384 well plates |  | M
M2 |  | Fluorescent intensity measurement Spectral range (minimum): 350-700 nm (excitation) (minimum) 350-700 nm (emission) Guideline performance - please detail the specific perforamnce of your equipment Wavelength accuracy: ± 0.5 nm Wavelength reproducibility: ± 0.5 nm |  | M
M3 |  | Fluorescent intensity measurement Please state the built-in filters for your equipment and also detail any additional options available (for additional options please ensure these are present in the pricing schedule as options). |  | M
D1 |  | Fluorescent intensity measurement Adjustable bandwidth: 5-50 nm |  | D
M4 |  | Luminescent intensity measurement Wavelength range: 350-700nm Detection limit: 3 x 10-21 moles of luciferase (or better) Include standard filters (approx 5 - please define details of these) |  | M
M5 |  | UV-Vis absorbance Light source - e.g. xenon flash lamp (please state if your light source is difference and the spectral range is supports) Spectral range (minimum) 200-600 nm (please detail any capability that exceeds this range) Guideline performance - please detail the specific performance of your equipment Wavelength accuracy: ≤ 0.3 nm Wavelength reproducibility: ≤ 0.3 nm Precision @ 260 nm: < 0.2 % Accuracy @ 260 nm: < 0.5 % Limit of detection (nucleic acids): ≤ 1 ng/μl |  | M
M6 |  | Filtered luminescence measurement Spectral range: 370-700 nm Approximate limits of detection (ATP): - Luminescence (glow) ≤ 225 amol/well - Luminescence (flash) ≤ 12 amol/well Multi-colour luminescence - please define filters supplied, including any OD attenuation filters |  | M
M7 |  | BRET and FRET capability - please define the capability for your equipment. |  | M
M8 |  | Footprint: must fit onto a standard depth lab bench (max depth approx 700 mm) |  | M
M9 | Temperature Control | Temperature Range: Ambient +5 to 42 °C |  | M
O1 |  | Additional temperature control: if you are able to supply additional temperature range (either above or below the range stated above), or methods for temperature control please detail those here and ensure these are costed within the pricing schedule as options. |  | O
M10 | Cell Counting/viability | Cell counting/viability capability (integration into plate reader desirable) Size range approx.: 4-90 μm Please define full capability including (but not limited to): accuracy, reproducibility, cell concentration range, imaging speed |  | M
M11 | Cell Imaging | Cell imaging capability (integration into plate reader desirable) Please define full capability including (but not limited to): illumination source, image type, objective magnification, use of autofocus, resolution, read speed. |  | M
M12 | Control System | Define how your system is controlled including details of software requirements, compatibility requirements, plus any additional control information. |  | M
M13 |  | Please state whether hardware (i.e. laptop/desktop/tablet) to control the plate reader is included within your equipment proposal or if an additional piece of hardware is required to host the control software. Note that it is the response to this requirement that is mandatory and not the actual provision of hardware. If control hardware is available please include this as an option in the pricing schedule. |  | M
Installation & Utilities
M14 | Installation | Define standard installation procedure (i.e. are vendor engineers required on-site, commissioning requirements, etc.). |  | M
M15 | Site requirements | A full site requirements guide must be provided as part of the Tender response, including any requirements for specialist lifting and manual handling equipment |  | M
M16 | Electrical Power | The system should operate from a standard UK mains supply (240Vac, 50Hz, single phase) |  | M
M17 | Other Utilities | Please state any additional utilities required including 3 phase power supply, external supply of: cooling water, gases or cryogenic gases. Note that it is the response to this requirement that is mandatory rather than a specific capbility. |  | M
Software and Connectivity
M18 | Connectivity | Must either be directly networkable to the local CPI network or networkable via the control hardware (laptop/desktop/tablet) |  | M
Regulatory, Support & Warranties
M19 | CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided |  | M
M20 | Training | Training in basic use of equipment and software - please detail how/when this will be provided. |  | M
O2 | Training | Please detail any other training that is available (please ensure details are costed within the pricing schedule). |  | O
M21 | Warranties | Warranty period should cover at least 12 months from installation |  | M
O2 | Extended warranty | Extended warranty of up to 5 years (please detail length of extended warranty and include costs in the pricing schedule) |  | O
M22 | Service support | Clear definition of servicing/calibration requirements including whether this can be carried out on site or requires return of equipment |  | M
OFFICE USE ONLY - This section below this line is for CPI Evaluation and Internal USE only
| Notes for Evaluation Team: |  | Name of Bidder to be inserted here | 0 | Multimode plate reader (for research use only)
Please assess all of the answers given, with particular attention to the Mandatory answers provided, taking into account, the Desirable and Optional answers too, however, you do not need to score each individual answer, but based on the scoring range 0-5, just give one singluar score (default set at "5" in the red and yellow box below, and as we have to provide feedback to bidders, you must also provide comments on which section(s) above where you felt that they did not meet the requirements.
| Evaluation Panel |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com