## Drop Codes (hidden)
RoomClassification | IA | QualificationTest
Not Classified | Direct Impact | Yes
Grade A | Indirect Impact | No
Grade B | No Impact | n/a
Grade C
Grade D

## URS Document
User Requirement Specification |  |  | CPI Internal Tender Reference & Title:- |  | Microfocus X-ray source
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 35%
USER REQUIREMENT |  |  |  |  | Tenderer Response
Requirement codes |  |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as Desirable) |  |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  | Requirement |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
M1 | X-ray source | (Nominal) Focal spot size at 4W ≈ 5 μm (up to 7 μm acceptable) |  | M
M2 |  | Spot size at maximum power ≈ 50 μm |  | M
M3 |  | Power: Minimum 65 W, preferred 75 W maximum power output |  | M
M4 |  | Integrated high voltage power supply so that no high voltage cable connection is required |  | M
M5 |  | X-ray tube voltage operational range 40 kV up to 150 kV |  | M
M6 |  | X-ray tube current range 10 to 500 μA |  | M
M7 |  | X-ray beam angle < 50° |  | M
M8 | X-ray source control unit | Single phase, compatible with UK mains power (240V AC, 50Hz) |  | M
M9 |  | Power consumption < 220 W |  | M
M10 | General | Operating and storage humidity 20 to 85 % |  | M
M11 |  | Minimum operating (storage) ambient temperature: +10°C to +40°C (0°C to +50°C) |  | M
M12 |  | Standard interface: e.g. RS-232C (9 pin D sub connector) |  | M
M13 |  | Regulation and standards: RoHS: EN 50581 Cat 9 EMC: IEC/EN 61326-1 Safety: IEC/EN 61010-1 |  | M
M14 |  | Connectors for, and interoperability with, door interlock system |  | M
Installation & Utilities
M15 | Installation | Any specialist installation requirements must be defined as part of the tender response. |  | M
M16 | Site requirements | All weights and dimensions of equipment in the form they will be delivered in must be provided prior to delivery. |  | M
D1 | Other Utilities | The system will not require a 3 phase supply, cooling water, gases or cryogenic gases |  | D
Software and Connectivity
D2 | Connectivity | Please define connectivity requirements: including interface, any software required, plus any other connectivity requirements. |  | D
Regulatory, Support & Warranties
M17 | CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided |  | M
D3 | Training | Details of any training provision should be provided along with details of access to technical support |  | D
M18 | Warranties | Warranty period should cover at least 12 months from installation |  | M
O1 | Extended warranty | Extended warranty of up to 5 years (if this is available please ensure there is a related line in the pricing document) |  | O
D4 | Service support | Details of preventative maintenance programmes required (if any) above Warranty requirements |  | D
OFFICE USE ONLY - This section below this line is for CPI Evaluation and Internal USE only
| Notes for Evaluation Team: |  | Name of Bidder to be inserted here | 0 | Microfocus X-ray source
Please assess all of the answers given, with particular attention to the Mandatory answers provided, taking into account, the Desirable and Optional answers too, however, you do not need to score each individual answer, but based on the scoring range 0-5, just give one singluar score (default set at "5" in the red and yellow box below, and as we have to provide feedback to bidders, you must also provide comments on which section(s) above where you felt that they did not meet the requirements.
| Evaluation Panel |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com