Date:08 September 2021

Attendees: 
Sebastian Ambaye (BEIS-HMG), Simon Lancaster (BEIS-HMG), Don McLellan (BEIS-HMG), Caroline Acton (Met Office-HMG), Barry Roberts (Ultra), Nick Kyprianou (Ultra)

Purpose of the Meeting

To introduce representatives of both parties and ensure a common understanding of the nuclear monitoring services provided by Ultra Electronics for BEIS.

Background

A contract was placed in 1990’s for the “fixed monitors” where there a total of 93 monitors deployed in the field.

This contract was varied in 2014  to include the provision of “mobile monitors” (R3M), which total of 104 in number,4 and varied once again in 2016 to include “spectroscopic monitors”, a total of 10 in 2016 plus spare parts such as probes.

The service provided by Ultra covers planned maintenance, reactive maintenance, parts and storage of parts and monitoring.

# Contracted Service

# Core Services

The Core Services comprise of 93 gamma radiation detectors distributed geographically evenly across the UK. Ultra designed and built the monitors, referred to as ‘fixed monitors’. The circuitry for the ‘fixed monitor’ was an established product that Ultra had sold to other customers before the UK government purchased the ‘fixed monitor’ as part of this contract. Therefore, Ultra claim the Intellectual Property of this circuitry as pre-existing IPR.

The Core Services are operations and maintenance (O&M) activities which are broadly divided between planned and unplanned tasks.

Planned O&M

- Calibration of the monitor probe

- Monitoring of the network

## Calibration of monitor probe

Each ‘fixed monitor’ enclosure (Figure 1) has a one-year calibration cycle based on the requirements set by the UK government in the 1990s when the capability was deployed. The calibration of the monitor probe includes the following activities.

- Assessment, repair, and calibration of monitors deployed in the field, this takes place at Ultra’s facility in Wimborne.

- Where necessary, swap out of the monitor with a recently calibrated probe.

- Set up and testing of the monitor, so the data is integrated into the RIMNET system.

The calibration of the monitor is a critical part of the Ultra service offering. The monitor is calibrated based on an agreed standard to provide quality assurance that the monitor is recording the correct radiation levels. This capability requires Ultra to have a facility that enables radiation testing of these devices.

## Monitoring of the network

Ultra has a sub-contract in place with BT to monitor the communications network. This capability ensures the health of the network and identifies faults and issues at the earliest opportunity.

Ultra provides a Service Delivery Manager to ensure the smooth delivery of the contracted services.

Figure 1: Describes the end to end managed service delivered by Ultra

Figure 1 describes the end to end managed serviced delivered by Ultra. The BT Cardway network transmits dose rate data over a private MPLS network, routing to RIMNET Data Centres (Exeter and Waterton). The process enables the seamless integration of the monitoring data into the RIMNET system.

# Unplanned O&M

Unplanned O&M activities describe the reactive approach Ultra take to managing the performance of the network; this includes,

- Service desk, first-line support

- Logistics and field engineering team

## Service desk, first-line support

A 24-hour service desk provides first-line support for all fault calls. A fault is identified by logging a call with Ultra via their helpdesk or through pro-active network monitoring. Once a fault is identified, Ultra coordinate the response, including liaising with third-party providers such as BT to resolve the issue (second-and third-line support). If the fault is associated with the asset, a field engineer will be deployed to replace the “fixed monitor” and return the faulty asset to Ultra for repair.

## Logistics and field engineering team

Spares are distributed across five locations to enable full UK coverage by the field engineering team to meet SLAs of the contract. The field engineering support is outsourced to Siemens, with Ultra training the engineers to the standards required to support their products. The spares are maintained and calibrated, ready to be deployed in response to a fault on the network. It is estimated that 10% of the asset inventory is held as spares.

## Non-Core Services

The Non-Core Services are additional components that have been added to the existing Core contract and comprise of the following elements,

- Mobile R3M network

- Spectroscopic monitors

- Consultancy

## Mobile R3M network

The R3M network comprises 104 gamma monitors designed as a low-cost supplementary network to fill the geographical gaps in the fixed gamma monitor network. The R3M is a self-contained unit designed and manufactured by Ultra Electronics Ltd who claim Intellectual Property related to this product.

Figure 2 provides a simple diagram of the critical comms components of the R3M network

The monitor is available in two versions:

- Fixed – includes a basic Monitor and mounting kit and is intended for installation on a fence-line, a post, or a wall, where an external power supply is available.

- Mobile – the basic monitor, is fitted inside a protective ‘traffic cone’ cover to reduce unwanted attention in public areas. This version is free-standing and can be deployed as required for an incident or other monitoring purposes.

Figure 2 describes the various components of the R3M network.

## R3M Operations and Maintenance

The Met Office is the primary contractor for the R3M O&M activities, including help desk and field engineering support. The Met Office has set up maintenance arrangements with Ultra for the following services to manage contract performance,

- Calibration of the monitor probe

- Repair of the faulty Monitor

- Spares

- Communications charges for SIM Cards and SMTP network services

The R3M has a three-year calibration cycle that includes the following activities,

- The Met Office engineer will remove the existing monitor for assessment, repair, and calibration and send it to Ultra’s factory in Wimborne.

- Ultra will assess, repair, and recalibrate the R3M based on the deployment schedule agreed with the Met Office.

- The Met Office will swap out of the monitor with a recently calibrated probe.

- The Met Office will set up and test the monitor to integrate the data into the RIMNET system.

For unplanned O&M activities, Ultra maintain and manage spares that are provided to the Met Office. There is no SLA for R3M availability (i.e., if a monitor fails in the field, how soon should it be fixed and the data accessible). The variations in site arrangements and monitor ownership means it would be difficult to set up a single SLA that applies a consistent availability requirement for all R3M monitors.

## Spectroscopic monitors

In September 2016, the Met Office deployed 10 Spectroscopic monitors at a Met Office site in Camborne, Cornwall. The monitors are supported under the existing O&M contract with Ultra.

## Consultancy

Ultra focus on the design and supply of safety-critical sensors and control systems mainly to the nuclear industry in the UK, North America, and China. Their expertise and experience have been provided to BEIS as an ad-hoc consultancy.

## Supply Chain

Ultra subcontract large elements of the services to Siemens, who provide field engineering support and have been working with Ultra for a significant amount of time. In addition, as mentioned previously, BT are a sub-contractor for communications-related elements of the services.

Summary

- Ultra Electronics intellectual property relates to the circuitry of the fixed and R3M monitors.

- The R3M monitors were co-designed by Met Office and Ultra.

- Both the fixed and R3M monitors have been purchased by BEIS and the department can do with them as they wish, this includes a ‘right to repair’. On theses assets can the following be provided:
(a)  Service history of assets

(b)	Confirmation that the only item in scope of IPR is the circuit board, can the     specific IP be shared i.e is this a patent or trade secret, can this be explained?

- It was mentioned in the meeting that it could be reverse engineered, but would come with complexity and risk. Can this be explained?

- We understand that Ultra would not have any concerns from a legal or IP issue prospective should the service currently provided by Ultra be tendered to a third party?

- The spectroscopic monitors have been purchased from a third-party (Would you clarify the supplier please?), and the design of these is to be considered proprietary IP of that organisation.