Schedule A

### THE HEALTH AND SAFETY EXECUTIVE

Tender Ref T34 1516

STATEMENT OF REQUIREMENTS FOR THE PROVISION OF SUPPLY OF A

SPARK Generator FOR PRODUCTION OF nANO pARTICLES

1	THE HEALTH AND SAFETY EXECUTIVE (HSE)

1.1	The Health and Safety Executive (HSE) is a Crown non-departmental public body with specific statutory functions in relation to health and safety.  It is appointed by the Secretary of State for Work and Pensions and employs around 3500 staff including policy advisers, inspectors, technologists and scientific and medical advisers. HSE’s job is to prevent people being killed, injured or made ill by work.

  - HSE consists of a governing Board comprising of a Chair and nine non-executive members, all of whom are appointed by the Secretary of State for Work and Pensions after consultation with organisations representing employers, employees, local authorities and others, as appropriate.

1.3	HSE comprises various directorates and groups and is led by a senior management team.  The Health and Safety Laboratory (HSL) is part of HSE.  HSE works from over 30 locations throughout Great Britain.

2.	Health & Safety Laboratory

The Health and Safety Laboratory (HSL) is Britain's leading industrial health and safety facility. HSL supports HSE’s mission to protect people's health and safety by ensuring risks in the changing workplace are properly controlled.

3	BACKGROUND TO THE PROJECT

3.1	For the calibration/validation of air sampling methods, sampler devices, sensors and the production of laboratory proficiency test samples, HSL requires the ability to produce stable standard particle atmospheres.

4		OBJECTIVES

4.1.	Production of stable standard particle atmospheres with outputs of constant size and concentration.

- SPECIFICATION REQUIRED

The requirements are for:-

- Capability to produce carbonaceous nanoparticles in the range 20 nm – 200 nm based upon a spark discharge process with the ability to produce other metallic nanoparticles.

- Mass flow of up to 10 mg/hour

- Ability to produce nanoparticle output for at least 5 hours stable

- Volume dilution output from generator (vendor to describe available options) ideally through use of mass flow controllers for precise dilution control

Installation, Commissioning and Training.  Tender to include itemised delivery, installation, and commissioning costs for the HSL Buxton site and provision of on-site training of 3-4 staff in the use of the new system.

Warranty.  Customary warranty or extended warranty to be offered. HSL would like to a achieve 24 month warranty period as standard for the instrument.

Other Services.

We would also like to look at preventative maintenance and calibration services