# Lot A: Specification of A New High-resolution Transmission Electron Cryo-microscope at MRC’s Laboratory of Molecular Biology, Cambridge

# The supplier of a new high-resolution transmission electron cryo-microscope will be responsible for the following three sections of the requirement:

- One electron microscope with 80 to 300 keV electron energy for cryoEM and electron cryo-tomography (cryoET);

- An extra 2 years of warranty;

- Moving and re-commissioning of an existing used Polara electron microscope (Polara I).

# 1.0.	Specification of an electron cryo-microscope with 80 to 300 keV electron energy for cryoEM and cryoET

# Summary of microscope deliverables:

  - To supply a complete high-resolution transmission Cryo-TEM

  - The microscope is expected to fulfil the specification as a high-end instrument capable of data collection by means of single particle EM work (cryoEM) as well as electron cryo-tomography work (cryoET)

  - To have automatic specimen exchange system

  - To have a phase plate

  - To have an energy filter for tomography

  - To have direct electron detectors for data collection with and without  energy-filtering

  - To have an assistant camera

  - To be user-friendly and have ability to run automated data acquisition

  - To have customer-witness acceptance test in the factory before delivery and to be delivered between December 2017 and January 2018

  - To have expected on-site acceptance before 1st April 2018

  - To be adapted to the physical room constraint

  - Services & Support

  - Innovation & added value are desirable

# Specification

  - A complete high-resolution transmission Cryo-TEM

    - To provide a complete system.

    - To accept full responsibility for provision and installation of a complete and fully operational system if third party suppliers are involved.

  - The core microscope should have the following specifications.

# Electron Beam Generation:

    - The generated electron beam should have electron energy of 80 to 300 keV (acceleration voltage). LMB expects but not limits to the following specifications.

      - It should be possible to adjust easily and in a few minutes the voltage to any voltage between 80 keV and 300 keV.

      - To have factory alignments at 300, 200, 120, and 80 keV.

      - 1.2.1.3 To be fitted with a field emission electron gun (FEG) to generate an electron beam with the highest possible brightness to provide enhanced spatial coherence, and with high emission stability for high-resolution imaging.

      - Energy spread of the electron beam to be ≤ 1.0 eV.

      - Beam spot size  1 nm and probe current ≥ 2 nA. The intensity of the beam should not decrease in time and should be constant for at least 2 days.

      - Lifetime of FEG tip to be > 2 years.

# Lens design:

    - Stable electron-magnetic lens design with constant power supply during a normal range of operations to provide maximum stability and reproducibility by minimizing thermal drift, lens hysteresis, cross-talk between optical components and to have fast switching between different operational modes.

# Condenser lens system:

    - Capable of producing parallel beam illumination at the specimen.

      - Ideally to have 3 condenser lenses, and with display of illumination geometry in the user interface.

      - Maximum spot drift 0.5 nm/min.

      - Quantitative indications of convergence angle, current on screen, size of illumination area, and electron dose in electrons per square Ångstrom per second (e-/Å2s) are desirable.

# Objective lens

    - To have a single condenser/objective lens with minimized aberrations at the eucentric position.

      - The spherical (Cs) and chromatic (Cc) aberration coefficients should be equal to or better than 2.7 mm at all energy levels.

      - To have a point resolution of better than 0.25 nm, and a line resolution of better than 0.14 nm (information limit).

# Specimen stage:

    - Must be fitted with a sample stage that is suitable for cryoEM and cryoET use

      - Temperature at sample must be < 85 K.

      - The holding time of the liquid nitrogen reservoirs should be more than 96 hours, or an automatic filling system should be provided.

      - Should be computer controlled with a range of movement specified in μm for the X, Y and Z axis and in degrees for alpha tilt.

      - Should have high stability: maximum stage drift rate < 0.1 Å / s at cryogenic temperature.

      - Should have reproducibility of ≤ 0.5 µm after specimen shift for 500 µm in X or Y direction.

    - The specimen height must be adjustable to allow eucentric tilting.

# 1.2.6.1   The α-tilt range should at least -70° to +70°.

# 1.2.6.2   The eucentricity during ± 70o tilting should be ≤ 2 µm in X and Y, and ≤ 4 µm in Z (defocus change).

# 1.2.6.3   To be supplied with a stage tracking system, capable of recording and re-calling the specimen areas, in digital and graphic forms.

    - The ice contamination rate must be minimised in microscope column.

# 1.2.7.1  The ice contamination rate must be < 1 Å / hour.

# 1.2.7.2  An automatic drift compensation system is desirable.

# Magnification:

# 1.2.8 The microscope should be fitted with intermediate & projector lenses to provide a highly reproducible, rotation-reduced, minimised anisotropic magnification in a magnification range with the following minimum specifications.

# 1.2.8.1 To produce a magnification range from 50x to 500,000x or higher.

# 1.2.8.2 Camera lengths in diffraction mode should be 200 mm (or less) to 5000 mm (or greater) at 300 keV.

# 1.2.8.3 Magnification should be reproducible to within ± 1.5%.

# 1.2.8.4 Should minimise the degree of rotation as magnification changes.

# 1.2.8.5 Anisotropic magnification distortion should be less than 0.1% at magnification above10,000x.

# Apertures:

# 1.2.9 The microscope aperture system must conform the following requirements.

# 1.2.9.1 To have an objective aperture holder with at least four aperture positions appropriate for different imaging conditions and phase plate.

# 1.2.9.2 To have two condenser aperture holders (C1 and C2), each with at least four apertures in each holder.

# 1.2.9.3 To have a selected area (SA) aperture holder, with at least four apertures.

# 1.2.9.4 All aperture holders must be motorized to maximize the degree of automation.

# Vacuum system:

# 1.2.10 To be expected to operate at ultra-high vacuum to both minimise sample contamination and maximise the life of the electron source, especially for cryo-EM use.

# 1.2.10.1 Pressures in the specimen chamber should be < 1 x 10-6 Pa.  Pressures in the gun chamber should be < 1 x 10-7 Pa.

# 1.2.10.2 Oil-free vacuum is required for the microscope and specimen exchange systems.

# 1.2.10.3 Fully interlocked and differential pumping is expected.

# 1.2.10.4 A specimen anti-contamination device must be provided and this should include a Dewar capable of achieving a cryogenic ‘hold time’ of at least 96 hours.

# 1.3. Sample Exchanger

# 1.3.1 Must have an automatic specimen exchange system.

# 1.3.2 At least 8 grids per one loading of the sample changer without breaking of vacuum.

# 1.3.3 Automatic cooling of stored samples to liquid nitrogen temperature (< 85 K); to maintain vitreous ice for at least 72 hours.

# 1.3.4. Contamination free: ice contamination rate should be less than 1 Å / hour within the sample changer.

# 1.4 Phase Plate

# 1.4.1 It is expected to be fitted with a phase plate to improve the imaging of weak phase objects, such as frozen hydrated biological specimens, particularly the sections or lamellae of cells.

# 1.4.2 The phase plate should be robust with an average life time of at least 9 months under normal TEM operating conditions.

# 1.4.3 Compatible with automated data collection.

# 1.4.4 Ideally, phase shift should be stable over at least an entire tomography data collection (2 -3 hours).

# 1.5 Energy Filter (EF) for Tomography

# 1.5.1 To be equipped with an energy-filter suitable for zero-loss imaging.

# 1.5.2 The EF must be proposed to include the following key criteria.

# 1.5.2.1 Energy resolution requirements are better than 10 eV.

# Geometrical distortion should be maximal 1%, ideally much lower.

# 1.5.2.2 Chromatic aberration < 1 pixel / 10 eV.

# 1.5.2.3 Combined cryo-EM and EF instrument operating in zero energy-loss    mode: non-isochromaticity < 2 eV.

# 1.5.2.4 Ease of use is important, alignment and tuning of the energy-filter is  expected to be as automatic as possible.

# 1.5.2.5 Sensitivity to magnetic / electric fields is minimal / compensated.

# 1.5.2.6 Filtered images should be recorded on a direct electron detector.

# 1.5.2.7 A second, fast buddy camera capable of very high dose rates would be useful.

# 1.5.2.8 Practical resolution requirements: on 4k x 4k images, Thon rings to Nyquist from amorphous carbon or Pt-Ir specimen at an exposure of less than 100 electrons per pixel.

# 1.5.2.9 To have alignment of EF at 300keV, 200keV and 80keV.

# 1.5.2.10 To have software for automatic filter alignment and zero-loss peak with distortion and aberration correction.

# 1.5.1.11 EF should have EFTEM spectrum imaging capability – with chosen slit width and energy range.

# 1.6 Direct Electron Detector(s)

# 1.6.1 The DEDs must have Direct Electron Detector(s) and be possible to record images with and without energy filtering.

    - The DEDs must meet or exceed the following key criteria.

      - Must have a minimum of 4k x 4k pixels.

      - To be highly sensitive for low-dose imaging (i.e. low read-out noise).

      - To have counting mode for images with and without energy filter.

      - To have software to collect data in summed and dose fractionated modalities.

      - To have a full frame rate > 300 fps (frame per second) for at least one detector.

      - To have dose rate of up to 4 electrons / pixel / s (electrons per pixel per second) useful in counting for at least one detector.

      - To have data collection and real-time processing in integration, counting and super resolution modes for at least one detector.

      - Radiation hardness must enable a minimum of 1000 images per day for a year.

      - DQE should be excellent at low resolution, i.e. > 0.8 (300 keV, counting mode).

      - DQE should be better than 0.5 at 0.5 Nyquist (300 keV, counting mode).

      - The DEDs should be possible to display a noise whitened FFT diffraction pattern immediately after recording an image, especially in integration mode.

    - It is highly desirable that there is a commitment to any near-future upgrade of the detectors in line with advances in detector technology & techniques. Such detector upgrade(s) should not incur extra expenditure by the purchasing organisation for new/modified hardware or software at least within two years.

# 1.7 Assistant camera

# 1.7.1 An additional fast, general purpose CCD camera is expected to assist the high DQE detectors for alignment, low-dose search, focus and medium-resolution image recording, as well as electron diffraction.

# 1.7.2 The assistant camera should be fast, at least 1Kx1K pixels but preferably larger, have a frame rate of at least 15 frames per second at a resolution of at least 500 x 500, and the ability to display and update the power spectrum at the same rate.

# 1.7.3 The assistant camera should be radiation hardened.

# 1.8 System, user interface, software and automatic data collection

# 1.8.1 Remote Operation is expected including hardware and software.

# 1.8.2 The TEM offered in the submitted documentation will be expected to be intuitive & user-friendly, robust & adaptable with the capability to accept future upgrades & improvements.

# 1.8.2 To be fitted with controls that are simple enough to be understood by ‘non-expert’ users and robust enough for heavy daily use, usually for 24/7 use.

# 1.8.3 To have a user-friendly interface for different levels of users, will be clearly labelled and intuitive enough for ‘non-expert’ users.

# 1.8.4 Advanced controls to be hidden so that use is limited to expert users or service engineers. A software design to allow different level of users in terms of control and safety is desirable.

# 1.8.5 To have software to run automated or semi-automatic alignment of the microscope, phase plate, detectors and energy-filter.

# 1.8.6 To have low dose imaging capability for cryoEM.

# 1.8.7 To be able to run automated data collection for cryoEM and cryoET.

# 1.8.8 All detectors, phase plate, and energy-filter should be embedded with the main microscope operation software.

# 1.8.9 Expected to have high degree of automation and all user functionality should be software controlled.

# 1.8.10 Safety controls must be implemented within software as well as in hardware for protecting the operators, instrument, detectors and specimens.

# 1.8.11 Software for tomography data processing including alignment, 3D reconstruction and visualization is desirable.

# 1.8.12 To ensure that all microscope functionality required for fully automated single-particle and tomography data collection is made available for control by software including SerialEM (An automatic tomography data acquisition software developed by University of Colorado).

# 1.9 Customer-witnessed acceptance test and delivery.

# 1.9.1 Customer-witnessed acceptance test in the factory is expected before delivery.

# 1.9.2 Delivery must be between December 2017 and January 2018.

# 1.10 Commissioning must to be finished and on-site acceptance accomplished before 1st April 2018.

# 1.11 EM Room size: the proposed microscope has to be adapted to the physical room constraints: the floor space is ~ 8 x 4 m, and ceiling height 4 m.

# 1.12 Service and support: supplier to propose the services to provide:

# 1.12.1 Response time for a service call: 24h / 48 h response time is expected (24 hours response by telephone or remote diagnosis or visit by engineer, and 48 hours response onsite by engineer visit or remote service)

# 1.12.2 Onsite training and support is expected

# 1.12.3 Application Support is desirable

# 1.12.4 Remote diagnosis is highly desirable

# 1.13 Innovation & Added Value

# 1.13.1 The buyer will give due consideration to any tender response, whether of standard manufacture or of an alternative design solution that provides and delivers innovation, good design, accounts for on-going product development and clearly demonstrates measurable subsequent benefits.

# 1.13.2 MRC-LMB encourages the bidders to put forward value added solutions. The specimen/EM-grid preparation apparatus and specimen holder/grid cartridges/frozen grid exchange accessories are particularly desired by MRC-LMB.

# 1.15.3 Additional sample vitrification capacity is needed so consideration should be given to the provision of suitable instrumentation as part of the response.

# 2.0. An extra 2 years of warranty

  - An extra two years extended warranty is expected and this is will be on top of the standard one year manufacturer’s warranty.

  - Labour and parts should be included in the warranty.

  - 24h / 48h response time is expected for service calls as part of the warranty.

  - Preventive maintenance service visits once per year as part of the warranty.

# 3.0. Moving and re-commissioning of an existing Polara electron microscope (LMB’s Polara I)

  - The contractor is expected to move an existing Polara I TEM from its existing room to an adjacent room in the same building on the same building level. The move has to be in September or early - mid October 2017 in order to vacate the room to be modified as the room for the new cryo-TEM tendered for.

  - To re-install and re-commission the same Polara I to full working order as soon as possible after the moving, ideally within 1 month. The configuration and specification of Polara I is attached as separate document.

  - The destination room for Polara I is 3 meters away from the current Polara I Room. The new Polara I room dimensions are 4.3 x 3.62 m and ceiling height 3.07 m. The moving path has two double-doors. The door width is 1.40 m (1.46 m if door frames removed) and door height is 2.65 m. The entrance of new Polara Room is 1.16 m width and height 2.71 m. The path is at the same floor level and with a ceiling height of 2.71 m. The diagram of the Polara Room and moving path is attached.