## Sheet1
| User Requirement Specification | CPI Tender Reference & Title:- | 1745 | High throughput potentiostat/galvanostat system
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column C, of which the total Quality element score of this sheet will equate to value in cell F6 | Scoring Criteria 0 = no response/answer not relevant to requirement 1 = limited response to requirements 2 = less than adequate, filfils some of the basic requirements 3 = adequate response, fulfils all basic requirements 4 = good response, fulfils all requirements to a high standard 5 = excellent response, fulfils requirements to an exceptionally high standard |  | The total percentage available which will contribute towards the Quality scoring element of this tender | 50%
USER REQUIREMENT |  |  |  | TENDERER RESPONSE
Requirement Codes |  |  |  | Guidance
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as desirable) |  |  | Tenderers must provide responses within the following columns; 1. Column F - please provide details of your equipment functionality in relation to CPI requirements specified in column E 2. Column G - please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response |  | Y
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  |  | N
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
# | Parameter | Requirement | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N) | N
System Performance
M1 | Key instrument capability | Potentiostat/galvanostat system, single unit or combined, to provide at least 50 total channels capable of EIS measurements. Portions of these to be supplied to a higher specification as detailed below. If more than one system required, please detail the combinations required to support this. | M |  |  | 1 | 0 | 0
M2 |  | Ability to independently gather and analyse data from all channels simultaneously. Control unit capable of managing multiple channels and interfacing with a robust and reliable software included | M |  |  | 2 | 0 | 0
M3 | At least 10 high frequency capable EIS channels, able to run 3 electrode measurements | At least 10 channels capable of EIS measurements across a range from greater than 5 MHz to less than or equal to 10 mHz | M |  |  | 3 | 0 | 0
O1 |  | From point M3, ability to add between 1 and 6 additional channels to provide a total of 11 - 16 channels capable of EIS measurements across a range from greater than 5 MHz to less than or equal to 10 mHz. Please provide pricing per additional channel and required cell cables. | O |  |  | 3 | 1 | 0
M4 |  | All channels able to be used for 3 electrode experiments | M |  |  | 4 | 1 | 0
M5 |  | All channels capable of supplying currents from greater than or equal to 10 A to less than or equal to 10 µA | M |  |  | 5 | 1 | 0
D1 |  | Voltage capability from 1-12 V | D |  |  | 5 | 1 | 1
M6 |  | Current resolution of 800 fA | M |  |  | 6 | 1 | 1
D2 |  | Floating mode, Analog filtering, calibration board | D |  |  | 6 | 1 | 2
M7 |  | Lowest voltage resolution of less than or equal to 1 µV | M |  |  | 7 | 1 | 2
M8 |  | Full stability control mode (good bandwiths) | M |  |  | 8 | 1 | 2
M9 | 8 high current channels for 2 electrode measurements | At least 8 channels capable of EIS measurements across a range from greater than or equal to 10 kHz to less than or equal to 10 mHz | M |  |  | 9 | 1 | 2
M10 |  | All channels able to be used for 2 electrode experiments | M |  |  | 10 | 1 | 2
M11 |  | All channels capable of supplying currents from greater than or equal to 10 A to less than or equal to 1 mA | M |  |  | 11 | 1 | 2
M12 |  | Voltage capability from 1-7 V | M |  |  | 12 | 1 | 2
M13 | 32 standard resolution channels for battery cycling, 2 electrode measurements | At least 32 channels capable of EIS measurements across a range from greater than or equal to 10 kHz to less than or equal to 10 mHz | M |  |  | 13 | 1 | 2
M14 |  | All channels able to be used for 2 electrode experiments | M |  |  | 14 | 1 | 2
M15 |  | All channels capable of supplying currents from greater than or equal to 1 A to less than or equal to 1 mA | M |  |  | 15 | 1 | 2
M16 |  | Voltage capability from 1-7 V | M |  |  | 16 | 1 | 2
O2 | Temperature monitoring and pressure capability | Capability to monitor cell temperature | O |  |  | 16 | 2 | 2
O3 |  | Please detail any options that would enable pressure testing on cells, individually or on a series of 8 cells. | O |  |  | 16 | 3 | 2
M17 | Spares/Consumables etc | System to include modules for testing multiple coin cells at once. Please provide racks for holding eg 2 x 8 cells in this fashion | M |  |  | 17 | 3 | 2
M18 |  | System to include 24 single coin cell holders 20 mm in diameter. | M |  |  | 18 | 3 | 2
M19 |  | System to include at least 50 wiring loops to facilitate connection to external cell setups (in line with number of channels supplied). Wires to be at least 1.5 m in length. | M |  |  | 19 | 3 | 2
O4 |  | Please provide pricing detail for any further consumables that may be of interest | O |  |  | 19 | 4 | 2
O5 |  | Please supply a quote for test cells that allow electrochemical testing of electrodes in both 2-electrode and 3- electrode configurations. Quantities between 1-16 in line with point M3 and O1. | O |  |  | 19 | 5 | 2
O6 |  | Please detail and include a quote for any cells which are capable of using a docking station rather than wires. | O |  |  | 19 | 6 | 2
0 | Installation and Utilities |  |  |  |  | 19 | 6 | 2
M20 | Installation | Installation must be performed by qualifed engineers at CPI NetPark facility | M |  |  | 20 | 6 | 2
M21 | Site Requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment | M |  |  | 21 | 6 | 2
D3 | Electrical Power | Where possible, the system should operate from a standard UK mains supply (240Vac, 50Hz, single phase). Please indicate if an alternate supply is required. | D |  |  | 21 | 6 | 3
0 | Software and Connectivity |  |  |  |  | 21 | 6 | 3
D4 | Connectivity | Computer with capability to run potentiostat system. Including sufficient RAM for operation of instrument, multiple ethernet ports to support components whilst being able to be networked, and included with 24 inch monitor. If cannot supply, please indicate ideal system specification. | D |  |  | 21 | 6 | 4
M22 |  | System equipped with computer Interface (e.g. ethernet, USB etc) | M |  |  | 22 | 6 | 4
M23 |  | Fully equipped software package for instrument control, setup of multiple experiments, data collection and data analysis. | M |  |  | 23 | 6 | 4
M24 |  | Software capable of quickly analysing cell parameters from multiple data sets. | M |  |  | 24 | 6 | 4
M25 |  | Software able to be installed on multiple PC's for offline data processing. | M |  |  | 25 | 6 | 4
D5 |  | Software capability to network and access externally to review data. | D |  |  | 25 | 6 | 5
0 | Regulatory, Support and Warranties |  |  |  |  | 25 | 6 | 5
M26 | CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided | M |  |  | 26 | 6 | 5
M27 | Training | Full training should be provided along with access to technical support | M |  |  | 27 | 6 | 5
M28 | Warranties | Warranty period should cover at least 18 months from installation | M |  |  | 28 | 6 | 5
M29 | Extended Warranty | Extended warranty of up to 5 years | M |  |  | 29 | 6 | 5
M30 | Service Support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. | M |  |  | 30 | 6 | 5
M31 |  | Details of preventative maintenance programmes required (if any) above Warranty requirements | M |  |  | 31 | 6 | 5
[6 empty] | 31 | 6 | 5
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com