Schedule A

STATEMENT OF SERVICE REQUIREMENTS FOR THE PROVISION OF A

REPLACEMENT ROBOTIC IMMUNOASSAY SYSTEM

TENDER HSE/SD/T8

- THE HEALTH AND SAFETY EXECUTIVE (HSE)

  - The Health and Safety Executive (HSE) is a Crown non-departmental public body with specific statutory functions in relation to health and safety.  It is appointed by the Secretary of State for Work and Pensions and employs around 2,600 staff including policy advisers, inspectors, technologists and scientific and medical advisers. HSE’s job is to prevent people being killed, injured or made ill by work.

  - HSE consists of a governing Board comprising of a Chair and ten non-executive members, all of whom are appointed by the Secretary of State for Work and Pensions after consultation with organisations representing employers, employees, local authorities and others, as appropriate.

- 1.3	HSE comprises various directorates and groups and is led by a senior management team.  The Health and Safety Laboratory (HSL), which is HSE’s Health and Safety Laboratory and Chemicals Regulation Directorate (CRD) are also part of HSE.  HSE works from 30 locations throughout Great Britain.

2	HEALTH AND SAFETY LABORATORY (HSE’S HEALTH AND SAFETY LABORATORY

  - The Health and Safety Laboratory (HSL) is Britain's leading industrial health and safety facility. HSL supports HSE’s mission to protect people's health and safety by ensuring risks in the changing workplace are properly controlled.

3.	BACKGROUND TO THE PROJECT

  - HSL is inviting tenders for the supply of robotic immunoassay system.

  - HSE’s allergen monitoring service currently uses a 8 year old Tecan Evo system, with fixed pipetting probes.  This is used primarily to measure by immunoassay low levels of allergens that have been extracted from filters after air sampling. The automated immunassay system is also used in the development of immunoassays. Overall the automation has a medium workload, and would be in use on a daily basis.

  - We are looking to replace the current Tecan Evo instrument.  We are looking for a system that can be used by staff who do not have considerable technical knowledge of the equipment and programming software, and will allow us to reduce maintenance expenditure in the future. We understand that the use of dispoable pipettes will incurr increased consumable exposure, but we have some methods where the ability to ensure zero carry-over between samples is necessary.

- SPECIFICATION

System

  - A dedicated, fully open colorimetric ELISA (96 well microtiter plates) automation system where all the activities within an ELISA are integrated from sampling to result calculation. Routinely samples are analysed in LP4 tubes (12mm diameter x 75mm in length]. Our usual assays are sandwich immunoassays using antibody coated plates and incorporating an amplification step, i.e. (1) sample/standard addition, (2) incubation,  (3) plate washing, (4) biotinylated antibody addition, (5) incubation, (6) plate washing, (7) amplification reagent addition, (8) incubation, (9) plate washing, (10) colour reagent addition, (11) incubation, (12) stop reagent addition and (13) absorbance reading. However, assays of different protocols/formats are also run.

Ease of use

  - Easy to operate for routine stored methods; software that guides an operator in placement of samples, reagents and wash solutions in appropriate positions, selection of tests for samples and setting up worksheets;

  - Easy to program new methods using a graphical interface and/or guided sequencing of necessary steps to develop new assay;

  - Has an automatic self test to ensure all functions of the instrument are working and desirable for available performance checks for the reader, washer, dilutors;

  - It is desirable that the acquired system allows for optimal time utilisation in running multiple batches of samples.

Capacity

  - A system capable of handling a run of up to four microtiter plates and handling up to about 90-100 samples in a run;

  - A system capable of handling at least 4 different assays within a batch; with our current way of working this means handling 4 different plates within a batch;

  - At least two different positions for sample diluents, 10 positions for reagents, 12 calibrator/QC position;

  - Ability to use two different buffers for washing microtiter plates.

Measurement & pipetting functions

  - Photometer capable at measuring accurately between 0-3.0 OD with an imprecision CV of less than 2%. Filters at 405 nm, 450 nm, 492 nm, 560nm, 620nm. Desirable to be capable of bichromatic readings;

  - Capable of adding a single reagent (100ul)  across a whole plate in less than 5 minutes;

  - Coefficient of precision of sample pipetting of <3% at 20 ul. It would be desirable if bidders could supply sample pipetting precision data for volumes in the range 10-100 ul;

  - Coefficient of precision of reagent pipetting of < 3% at 50 ul and above;

  - It would be advantageous for the system to be capable of using either disposable or fixed tips for sample/reagent additions, but a disposable tip system would meet our needs;

  - It would desirable for the system to be capable of serial dilutions of any sample in the range of 1:1 to 1:10.

Additional functionality

  - It is desirable that the instrument has an integral QC storage and plotting program

  - The availability of programmable, thermostated incubators for microtiter plates would be desirable;

  - An integral sample bar code reader would be desirable.

Updates & support

  - Software updates offered as they become available;

  - Servicing or responding to issues with the system should be carried out by UK based personnel able to respond within a 48 hour period.

- DELIVERY

  - Delivery is to be made to the Health & Safety Laboratory, Buxton, Derbyshire, SK17 9JN.  Please provide costs for delivery and anticipated lead delivery times.

  - It is desirable that delivery and installation can be completed within 6 weeks of notification of the successful tender.

- INSTALLATION, COMMISSIONING and TRADE-IN OF EXISTING MACHINE

  - The new machine is to be Installed and commissioned at HSE Buxton.

  - HSL is considering the possibility of a trade-in for the existing machine.  If this is of interest, please submit you offer within the pricing schedule (Schedule D).

- TRAINING

  - Provide training to six HSE staff on the new machine and software

- WARRANTY AND DOCUMENTATION

  - We would like to achieve an initial 24 month warranty period for the instrument.

  - All user manuals and copies of software should be received with the instrument. Any PC is considered part of the system we buy and should be specifically included any warranty agreement offered by the tenderer. We would expect that any problems with hardware and /or software that forms part of the system should be resolved by the supplier

- INFORMATION TECHNOLOGY.

  - HSL has a separate Facilities Management IT Provider who are responsible for the whole network of the laboratory.  In that regard, they have a clear responsibility for ‘impact testing’ new IT equipment onto HSL’s IT network. The laboratory’s FM IT Provider sets a minimum requirement on all new equipment.

  - It is therefore a requirement that all Tenderers should specify in the fullest detail possible, the makers name and specification etc of all IT equipment (hardware, software etc) that they propose to supply including details of the proposed operating system. Details of any associated IT peripherals and software should also be supplied. It would greatly assist if separate itemised costs for all of the associated IT components (e.g PC, LCD, Printer) could be included within the tender response, should these be available.

  - In addition, please provide warranty information, stating whether it is standard warranty (return to base) or onsite warranty (next business day), together with the length of warranty that will be supplied with the proposed equipment

  - Please also provide details of exactly how these warranties shall be activated and/or requested, should any problems occur.

- SOFTWARE

  - Prior to placing any order, HSL requires assurance from you that the proposed software contained within the order does not require administration rights to run and operate correctly post installation. This is because HSL is unable to run and operate software in MS Windows that needs to have administration rights before it will run and operate correctly post installation. Therefore would you:

- Confirm that the software does not require administration rights to run;  or

2) 	If you are unable to provide a version of software that does not run and operate correctly without admin rights, please advise what you can do to resolve this.

  - Note that in the event of an order being placed, and it is found that "administration rights" are required to run and operate the software, or, alternatively, that you do not notify HSL's Purchasing and Procurement Section that the software needs administration rights in order to run and operate correctly, HSL reserves the right to cancel the order and to place an order with another supplier and to seek a refund of any payments that have been made against the purchase order.