## Drop Codes (hidden)
RoomClassification | IA | QualificationTest
Not Classified | Direct Impact | Yes
Grade A | Indirect Impact | No
Grade B | No Impact | n/a
Grade C
Grade D

## URS Document
User Requirement Specification |  |  | CPI Internal Tender Reference & Title:- | Ref 1103 | Coxon Building –High Throughput Expansion
The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 25%
USER REQUIREMENT |  |  |  |  | Tenderer Response
Requirement codes |  |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as Desirable) |  |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  | Requirement |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
Building Works
Floor |  | Vendor to confirm concrete floor is level. Needs to be within 5mm over a 2m distance in any direction. Then screed / self level as neccesary. Vendor to describe how they will achive this. |  | M
Floor |  | Vendor to also provide a price for the floor level specification of a 2mm over a 2m distance in any direction. Then screed or use suitable self leveling compound as necessary. |  | O
Chemical waste drain |  | Vendor to design a chemical drain system to serve the relevant equipment, using the minimum of floor penetrations / core drilling. How will you achieve this? |  | M
Changing room / viewing area |  | Vendor to design, supply and install a suitable room in line with the specification, including stud partition walls, doors, windows, light fittings, underdrawn ceiling, floor covering etc. |  | M
Building Design and Instalation
Installation |  | Installation must be performed by qualifed / competent trades persons , such as electrical installers must be 18th edition trained , ATEX installers and ispectors must be COMPEX trained etc. Vendor must be able to supply training certs / CV's of all installation personnel |  | M
Lifting / Manual handling |  | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment |  | M
Air handling / LEV |  | Air handling and LEV systems must be designed by a competent person . Both systems must be compatible and ultimatly the end result will be a fully balanced system which has a 50% spare capacity . How will this be achieved? |  | M
Electrical Power / Lighting |  | Electrical installation must comply fully with the latest 18th edition wiring regulations and best practice. Cable calculations to be be supplied for the distribution board(s) and all equipment items. All 13A socket outlets both fixed and on portable benches must be RCD protected. |  | M
UPS (uninteruptable power supplies) |  | 1x Formax, 1x Swing and 1x Labman robots each require an individual UPS system. To be sited local to the robot. Vendor to confirm proposed system to achive the required specification. Battery outonomy of 15 minutes. |  | O
Planning & Building Regulations
Planning / Buiding Regulations |  | Vendor to apply for any required building regulations also comply with fire regulations |  | M
Regulatory, Support & Warranties
CE Marking |  | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided |  | M
Warranties |  | Warranty period should cover at least 18 months from installation / sign off. |  | M
Extended warranty |  | Vendor to quote for an additional 3 and 5 years maintenance / warranty package |  | O
Waste management |  | Vendor to supply copy of waste management certificate / procedure. |  | M
Health and safety |  | Specific risk assessments and method statements must be produced for all work activities. Also detailed programmes for lifting operations in conjunction with CPI and Facilities Management. CPI have a strict policy regarding health and safety / PPE . All contractors will require a site induction which should last no more than 1 hour. If the contract is notifiable under CDM regulations then then principle contractor must comply with all the requirements under this legislation. Vendor to supply examples of risk assessments and method statements appropriate for this type of propject. |  | M
OFFICE USE ONLY - This section below this line is for CPI Evaluation and Internal USE only
| Notes for Evaluation Team: |  | Name of Bidder to be inserted here | Ref 1103 | Coxon Building –High Throughput Expansion
Please assess all of the answers given, with particular attention to the Mandatory answers provided, taking into account, the Desirable and Optional answers too, however, you do not need to score each individual answer, but based on the scoring range 0-5, just give one singluar score (default set at "5" in the red and yellow box below, and as we have to provide feedback to bidders, you must also provide comments on which section(s) above where you felt that they did not meet the requirements.
| Evaluation Panel |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com