## F017 HTE Lab
F1.24 High Throughput Formulation Lab Equipment List
Key:
Added to Layout/
Future Equipment
[5 empty] | Dimensions (excluding PC and/or monitor) | [5 empty] | Services
Instrument | Asset No | Tender Type | Instruction | Code | Width (mm) | Depth (mm) | Height (mm) | Weight (kg) | Notes | Floor Standing or Bench | Power Supply | No Of Plugs | CA (bar) (Assume 6bar supply) | N2 (bar) (Assume 6bar supply) | Other Gas | Water | Solvent Extract Requirement | Solvent Extract connection size | Solvent Extract Velocity m3/hr | Powder Extract Requirement | Powder Extract connection size mm | Powder Extract Velocity m3/hr | Waste | Vacuum | Power Output (W) | No of data points
Formax Formulation Robot - comprising | FC00167 | Vendor to install supplies at Low level. | Existing Equipment Manufacturer to Move | [7 empty] | On wall 1x 3PH N/PE 400VAC/230V AC 50~60Hz 6000VA (Autoplant) 1x 3PH N/PE 400VAC/230V AC 50~60Hz 4x1080VA (Autoplant) 1x 16A Single Phase 100-240V AC 50~60Hz 1700VA (E Cabinet) 1x 16A Single Phase (Robot) With in line junction box 1x 240v 1~50Hz max 7.0A single plug (Unistat Unichiller 015-MPC plus) 1x 240v 1~50Hz max 13A single plug (Unistat Tango Unichiller) 1x 240v 1~50Hz max 13A single plug (Unistat Tango Unichiller) Supplies allocated On Ladder rack 1 1x 240v single plug (Computer) 1x 240v single plug (Monitor) 1x 240v single plug (Bruker Matrix F) 1x 240v single plug (PC 3000 series) 1x 240v single plug (Data router) 1x 240v single plug 1x 240v single plug 1x 240v single plug | REFERENCE POWER SUPPLY | YES TYPE 2 | YES TYPE 3 | N/A | N/A | SOLVENT Ref LEV082 | Reference: LEV082 | eference: LEV082 | N/A | N/A | N/A | N/A | YES STAND ALONE PUMP | N/A/ | 1x Double Data Points (Installed on Ladder Rack 1)
Formax cabinet |  |  |  |  | 2350 | 950 | ≤2000 | 750 |  | Floor. South wall
Electronics Cabinet |  |  |  |  | 610 | 900? |  | 100 |  | Floor. South wall
Module Control Box |  |  |  |  | 630 | 900 |  | 250 |  | Floor. South wall
Tango Nuevo x 2 |  |  |  |  | 350 | 500 |  | 56 |  | Floor. South wall
Huber Unichiller UC015 |  |  |  |  | 400 | 500 |  | 52 |  | Floor. South wall
Vacuum pump |  |  |  |  | 200 | 200 |  |  |  | Floor. South wall
NIR |  |  |  |  | 300 | 300 |  |  |  | Floor. South wall
PSD |  |  |  |  | 300 | 300 |  |  |  | Floor. South wall
NIR/PSD/Rheo PC |  |  |  |  | 300 | 300 |  |  |  | Floor. South wall
Formax extension with Rheometer |  |  |  |  | 950 | 890 | 1920 | 250 |  | Floor. South wall
Waste container / cabinet |  |  |  |  | 300 | 300 |  |  |  | Floor. South wall
Swing Formulation Robot - comprising | FC00169 | Vendor to install supplies at Low level. | Existing Equipment Manufacturer to Move | [7 empty] | On wall 1x 16A Single Phase 100-240V AC 50~60Hz 1700VA (E Cabinet) 1x 240v 1~50Hz max 13A single plug (Unistat Tango Unichiller) Supplies allocated On Ladder rack 1 1x 240v single plug (Vacuum pump) 1x 240v single plug (Robot control box) 1x 240v single plug (Robot control box) | REFERENCE POWER SUPPLY | YES TYPE 2 | YES TYPE 3 | N/A | N/A | SOLVENT Ref LEV083 | Reference: LEV083 | Reference: LEV083 | N/A | N/A | N/A | N/A | YES STAND ALONE PUMP | N/A | 1x Double Data Points. (Installed on Ladder Rack 1)
Swing Cabinet |  |  |  |  | 1480 | 910 | <2000 | 350 |  | Floor. South wall
Electronics Cabinet |  |  |  |  | 610 | 910 |  |  |  | Floor. South wall
Module Control Box |  |  |  |  | 600 | 800 | 1290 |  |  | Floor. South wall
Waste container / cabinet |  |  |  |  | 300 | 300 |  |  |  | Floor. South wall
Crystat |  |  |  |  | 400 | 500 |  | 52 |  | Floor. South wall
New Formulation Robot | Future | Vendor to install supplies at High level. | Future Purchase. |  | 2000 | 8000 |  |  |  | Floor. Central. | 1x 16A Single Phase | REFERENCE POWER SUPPLY | YES TYPE 2 | YES TYPE 3 | N/A | N/A | SOLVENT Ref: LEV091 | 4" | 0.4 SOLVENT | POWDER REF: LEV093 | 4" | x.x POWDER | N/A | YES STAND ALONE PUMP |  | 2x Double Data Points
Kruss LSA | FC00165 | Vendor to install supplies. (Supplies on Ladder rack 2). | Existing Equipment CPI to Move | KLSA | 1010 | 960 | 540 | 100 | Supplies allocated on Ladder rack 2. Solvent extract to be installed at high level. Extract boothe under design. On dedicated stand. | Dedicated stand. Ladder rack 2. | 2x Standard 240v switched sockets. | 4 | N/A | N/A | N/A | N/A | SOLVENT Ref LEV098 | SOLVENT Ref LEV098 | SOLVENT Ref LEV098 | N/A | N/A | N/A | N/A | N/A | N/A | 1x Double Data Points. (Installed on Ladder Rack 2)
Optical Coatings Measurement | FC00166 | Vendor to install supplies. (Supplies on Ladder rack 2). | Existing Equipment Manufacturer to Move | CPOC | 1600 | 1000 | 1800 | 500 | Supplies allocated on Ladder rack 2. | Floor. Ladder rack 2. | 2x Standard 240v switched sockets. | 3 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1x Double Data Points. (Installed on Ladder Rack 2)
Labman CIPD Coatings Application | FC00006 | Vendor to install supplies. (Supplies onwall dado trunking). | Existing Equipment Manufacturer to Move | CPID | 1500 | 1000 | 1800 | 500 |  | Floor. South wall. | 2x Standard 240v switched sockets. | 3 | YES TYPE 2 | N/A | N/A | N/A | SOLVENT Ref: LEV097 | Reference: LEV097 | Reference: LEV097 | N/A | N/A | N/A | N/A | N/A | N/A | 1x Double Data Points. (Supplies onwall dado trunking).
Labman CPRS Electrical Testing | FC00007 | Vendor to install supplies. (Supplies on wall dado trunking). | Existing Equipment Manufacturer to Move | CPRS | 1500 | 1000 | 1800 | 200 |  | Floor. South wall. | 2x Standard 240v switched sockets. | 3 | -- | -- | -- | -- | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1x Double Data Points. (Supplies onwall dado trunking).
Light Transmission | FC00150 | Vendor to install supplies. (Supplies on Ladder rack 2). | Existing Equipment Manufacturer to Move | CPLT | 1520 | 960 | 1140 | 300 | Supplies allocated on Ladder rack 2. | Floor. Ladder rack 2 | 1x Standard 240v switched sockets. | 2 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1x Double Data Points. (Installed on Ladder Rack 2)
[11 empty] | 2x Standard 240v switched sockets. | [11 empty] | --
FTIR Plate Reader | FC00160 | Vendor to install supplies. (Supplies on Bench 5). | Existing Equipment CPI to Move Bench 5 | PRIR | 1000 | 750 | 500 | 50 | Ideally in corner as instrument is L shaped | Bench 5 | 2x Standard 240v switched sockets. | 3 | N/A | N/A | N/A | N/A | SOLVENT Ref LEV099 | Reference: LEV099 | Reference: LEV099 | N/A | N/A | N/A | N/A | N/A | N/A | 1x Double Data Points. (Installed on Bench 5)
Turbiscan | FC00143 | Vendor to install supplies. (Supplies on Bench 4). | Existing Equipment CPI to Move Bench 4 | TURB | 950 | 750 | 850 | 50 | Axis means this requires > standard depth bench or overlap | Bench | 1x Standard 240v switched sockets. | 1 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1x Double Data Points. (Installed on Bench 4)
Lumisizer | FC00151 | Vendor to install supplies. (Supplies on Bench 5). | Existing Equipment CPI to Move Bench 5 | LUMI | 370 | 570 | 260 | 50 |  | Bench | 2x Standard 240v switched sockets. | 3 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1x Double Data Points. (Installed on Bench 5)
Cure Oven N2 | Future | Vendor to install supplies. (Supplies on North wall dado trunking). | Future. CPI to Purchase | Ovens | 750 | 750 | 750 | 40 |  | Bench | 1x Standard 240v switched sockets. | 1 | N/A | YES TYPE 2 | N/A | N/A | SOLVENT Ref LEV086 | Reference: LEV086 | Reference: LEV086 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Lab Drying oven | NEW | Vendor to install supplies. (Supplies on North wall dado trunking). | New CPI to Purchase. CPI to install. | DCAB | 1000 | 750 | 500 | 40 |  | Bench | 1x Standard 240v switched sockets. | 1 | N/A | N/A | N/A | N/A | SOLVENT Ref LEV085 | Reference: LEV085 | Reference: LEV085 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Vac Oven | NEW | Vendor to install supplies. (Supplies on North wall dado trunking + 1 under bench). | New CPI to Purchase. CPI to install. | Ovens | 750 | 750 | 750 | 40 | Requires hole cutting in bench for cables and tube. Install 1 240v socket under bench for pump. Extract pipe to be installed for pump for heat. | Bench | 2x Standard 240v switched sockets. | 1 UNDER BENCH 1 ABOVE BENCH | N/A | N/A | N/A | N/A | SOLVENT Ref LEV084 | Reference: LEV084 | Reference: LEV084 | N/A | N/A | N/A | N/A | Yes Stand alone pump | N/A | N/A
Cure Oven 1 | Future | Vendor to install supplies. (Supplies on North wall dado trunking). | Future. CPI to Purchase | Ovens | 750 | 750 | 750 | 40 |  | Bench | 1x Standard 240v switched sockets. | 1 | N/A | YES TYPE 2 | N/A | N/A | SOLVENT Ref LEV087 | Reference: LEV087 | Reference: LEV087 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Cure Oven 2 | Future | Vendor to install supplies. (Supplies on North wall dado trunking). | Future. CPI to Purchase | Ovens | 750 | 750 | 750 | 40 |  | Bench | 1x Standard 240v switched sockets. | 1 | N/A | YES TYPE 2 | N/A | N/A | SOLVENT Ref LEV088 | Reference: LEV088 | Reference: LEV088 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Laboratory Dishwasher/solvent wash station | Not Required |  | Not Required |  | 813 | 889 | 711 | 114 |  | Floor | 1x Standard 240v switched sockets. | 1 | Y | -- | -- | Cold Mains | SOLVENT | XX mm |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A
DAC Mixer | Existing | Vendor to install supplies. (Suppy from Bench or Dado trunking) | CPI to Locate from existing equipment | DAC | 275 | 300 | 480 |  |  | Bench | 1x Standard 240v switched sockets. | 1 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Elga water purification | NEW | Vendor to install supplies. (Supplies on Bench 10 dado trunking + 1 Water supply above bench). | New CPI to Purchase | ELGA | 540 | 360 | 590 |  |  | Bench | 1x Standard 240v switched sockets. | 1 | N/A | N/A | N/A | Cold Mains Above bench | N/A | N/A | N/A | N/A | N/A | N/A | Loose Pipe into Sink | N/A | N/A | N/A
Tall Fridge | NEW | Vendor to install supplies. (Supplies on North wall dado trunking). | New CPI to Purchase and insta;l. | N/A | 600 | 600 | 2000 |  |  | Floor | 1x Standard 240v switched sockets. | 2 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Tall Freezer | NEW | Vendor to install supplies. (Supplies on North wall dado trunking). | New CPI to Purchase and install. | N/A | 600 | 600 | 2000 |  |  | Floor | 1x Standard 240v switched sockets. | 2 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Powder / Solvent Dispenser 1 | NEW | Vendor to install supplies. (Supplies on East wall Bench 8). | New CPI to Purchase and install. | N/A | 1000 | 700 | 600 |  |  | Bench | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT Ref LEV095 | SOLVENT Ref LEV095 | 0.5 SOLVENT | POWDER REF: LEV094 | POWDER REF: LEV094 | POWDER REF: LEV094 | N/A | N/A |  | N/A
DSC - Autosampler, Cooling Unit, PC, Keyboard & Mouse, Monitor | FC00158 | Vendor to install supplies. (Supplies on South wall Bench 7). | Existing Equipment in F0.17 CPI to Move | N/A | 1200 | 750 | 500 |  |  | Bench | 3x Standard 240v switched sockets. 1x Standard 240v switched sockets under bench. 1x 16A Single Phase | Ref Power Supply | N/A | YES TYPE 2 Comes with regulator and filter | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1x Double Data Points
Auto TGA | FC00313 | Vendor to install supplies. (Supplies on South wall Bench 7). | Existing Equipment in F0.17 Manufacturer to Move | N/A | 1200 | 750 | 500 |  |  | Bench | 1x Standard 240v switched sockets.. 1x 16A Single Phase to be mounted on wall. | Ref Power Supply | YES | YES TYPE 2 Comes with regulator and filter | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1x Double Data Points
LEV080 | LEV080 | Vender to Supply & Install. | Vendor Install High Level | FC024 | N/A | N/A | N/A | N/A | With Damper |  | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT | 300 | 0.5 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
LEV081 | LEV081 | Vender to Supply & Install. | Vendor Install High Level | FC023 | N/A | N/A | N/A | N/A | With Damper |  | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT | 300 | 0.5 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
LEV082 | LEV082 | Vender to Supply & Install. | Vendor Install High Level | (FC00167) FORMULATION ROBOT | N/A | N/A | N/A | N/A | With Damper @ 2500mm from F.F.L. |  | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT | 100 | 0.4 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
LEV083 | LEV083 | Vender to Supply & Install. | Vendor Install High Level | (FC00169) LIQUIDS ROBOT | N/A | N/A | N/A | N/A | With Damper @ 2500mm from F.F.L. |  | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT | 100 | 0.2 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
LEV084 | LEV084 | Vender to Supply & Install. | Vendor Install High Level | Vac Oven | N/A | N/A | N/A | N/A | With Damper @ 2500mm from F.F.L. |  | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT | 75 | 0.4 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
LEV085 | LEV085 | Vender to Supply & Install. | Vendor Install High Level | Lab Drying Oven | N/A | N/A | N/A | N/A | With Damper @ 2500mm from F.F.L. |  | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT | 75 | 0.4 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
LEV086 | LEV086 | Vender to Supply & Install. | Vendor Install High Level | Cure Oven N2 | N/A | N/A | N/A | N/A | With Damper @ 2500mm from F.F.L. |  | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT | 75 | 0.4 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
LEV087 | LEV087 | Vender to Supply & Install. | Vendor Install High Level | Cure Oven 1 | N/A | N/A | N/A | N/A | With Damper @ 2500mm from F.F.L. |  | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT | 75 | 0.4 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
LEV088 | LEV088 | Vender to Supply & Install. | Vendor Install High Level | Cure Oven 2 | N/A | N/A | N/A | N/A | With Damper @ 2500mm from F.F.L. |  | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT | 75 | 0.4 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
LEV089 | LEV089 | Vender to Supply & Install. | Vendor Install High Level | Ladder Rack 3 | N/A | N/A | N/A | N/A | With Damper @ 2500mm from F.F.L. |  | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT | 100 | 0.5 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
LEV090 | LEV090 | Vender to Supply & Install. | Vendor Install High Level | Ladder Rack 3 | N/A | N/A | N/A | N/A | With Damper @ 2500mm from F.F.L. |  | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT | 100 | 0.5 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
LEV091 | LEV091 | Vender to Supply & Install. | Vendor Install High Level | New Formulation Robot | N/A | N/A | N/A | N/A | With Damper @ 2500mm from F.F.L. |  | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT | 100 | 0.5 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
LEV092 | LEV092 | Vender to Supply & Install. | Vendor Install High Level | WFC021 | N/A | N/A | N/A | N/A | With Damper |  | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT | 450 | 0.5 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
LEV093 | LEV093 | Vender to Supply & Install. | Vendor Install High Level | New Formulation Robot | N/A | N/A | N/A | N/A | With Damper @ 2500mm from F.F.L. |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | POWDER | 4" | x.x POWDER | N/A | N/A | N/A | N/A
LEV094 | LEV094 | Vender to Supply & Install. | Vendor Install Low Level | Dispensor 1 | N/A | N/A | N/A | N/A | With Damper @ 2500mm from F.F.L. |  | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT | 100 | 0.5 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
LEV095 | LEV095 | Vender to Supply & Install. | Vendor Install Low Level | Dispensor 1 | N/A | N/A | N/A | N/A | With Damper @ 2500mm from F.F.L. |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | POWDER | 4" | x.x POWDER | N/A | N/A | N/A | N/A
LEV096 | LEV096 | Vender to Supply & Install. | Vendor Install High Level | FC022 | N/A | N/A | N/A | N/A | With Damper |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | POWDER | 4" | x.x POWDER | N/A | N/A | N/A | N/A
LEV097 | LEV097 | Vender to Supply & Install. | Vendor Install High Level | (FC00006) Labman CIPD | N/A | N/A | N/A | N/A | With Damper @ 2500mm from F.F.L. |  | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT | 100 | 0.4 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
LEV098 | LEV098 | Vender to Supply & Install. | Vendor Install High Level | Kruss LSA (FC00165) | N/A | N/A | N/A | N/A | With Damper @ 2500mm from F.F.L. |  | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT | 100 | 0.4 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
LEV099 | LEV099 | Vender to Supply & Install. | Vendor Install Low Level | FTIR Plate Reader (FC00160) | N/A | N/A | N/A | N/A | With Damper @ 2500mm from F.F.L. |  | N/A | N/A | N/A | N/A | N/A | N/A | SOLVENT | 100 | 0.4 | N/A | N/A | N/A | N/A | N/A | N/A | N/A
FC022 | FC022 | Vender to Supply & Install with Services. | Vendor Install |  | 1500 | 900 | Full Height | - | Power supply option to be provide by Vendor. | Floor | 2x Standard 240v switched sockets. | 4 | YES TYPE 2 | YES TYPE 2 | N/A | YES COLD MAINS | N/A | N/A | N/A | POWDER | POWDER: Ref LEV096 | x.x POWDER | YES | N/A | N/A | 2x Double Data Points
FC023 (2x Vented Acid / Base) | FC023 | Vender to Supply & Install with Services. | Vendor Install |  | 1800 | 900 | Full Height | - | Power supply option to be provide by Vendor. | Floor | 2x Standard 240v switched sockets. | 4 | YES TYPE 2 | YES TYPE 2 | N/A | YES COLD MAINS | SOLVENT Ref LEV081 | SOLVENT Ref LEV081 | 0.5SOLVENT Ref LEV081 | N/A | N/A | N/A | YES | N/A | N/A | 2x Double Data Points
FC024 (2x Vented Acid / Base) | FC024 | Vender to Supply & Install with Services. | Vendor Install |  | 1800 | 900 | Full Height | - | Power supply option to be provide by Vendor. | Floor | 2x Standard 240v switched sockets. | 4 | YES TYPE 2 | YES TYPE 2 | N/A | YES COLD MAINS | SOLVENT Ref LEV080 | SOLVENT Ref LEV080 | SOLVENT Ref LEV080 | N/A | N/A | N/A | YES | N/A | N/A | 2x Double Data Points
WFC021 | WFC021 | Vender to Supply & Install with Services. | Vendor Install |  | 3000 | 1300 | Full Height |  | Power supply option to be provide by Vendor. | Floor | 2x Standard 240v switched sockets. | 4 | YES TYPE 2 | YES TYPE 2 | N/A | YES COLD MAINS | SOLVENT Ref LEV092 | SOLVENT Ref LEV092 | SOLVENT Ref LEV092 | N/A | N/A | N/A | YES | N/A | N/A | 2x Double Data Points
Single Fixed Bench 1 | NEW | Vendor to Supply Fixed Benches install supplies at Low level. | Vendor Install |  | 750 | 5832 | 900 |  | Back boards to back and both ends | Floor | 6x Standard 240v switched sockets. | 12 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 2x Double Data Points
Single Fixed Bench 2 | NEW | Vendor to Supply Fixed Benches install supplies at Low level. | Vendor Install |  | 750 | 1546 | 900 |  | Back boards to back and both ends | Floor | 2x Standard 240v switched sockets. | 4 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1x Double Data Points
Double Fixed Bench 3 | NEW | Vendor to Supply Fixed Benches install supplies at Low level. | Vendor Install |  | 1530 | 2075 | 900 |  | Back boards to both ends | Floor | 4x Standard 240v switched sockets. | 8 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 1x Double Data Points
Double Mobile Bench 4 | NEW | Vender to Supply Mobile Benches (Services required See Note) | Vendor Install |  | 1530 | 1900 | 900 |  | No Backboards required. (Services required to be installed on benches with 1 mains lead with commado plug this will be hard wired into the RCD to protect the mains socket outlets. Plus data cables) | Floor | 4x Standard 240v switched sockets. | 8 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 2x Double Data Points
Mobile Double Bench 5 | NEW | Vender to Supply Mobile Benches (Services required See Note) | Vendor Install |  | 1530 | 1900 | 900 |  | No Backboards required. (Services required to be installed on benches with 1 mains lead with commado plug this will be hard wired into the RCD to protect the mains socket outlets. Plus data cables) | Floor | 4x Standard 240v switched sockets. | 8 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 2x Double Data Points
Mobile Single Bench 6 | NEW | Vender to Supply Mobile Benches (Services required See Note) | Vendor Install |  | 1530 | 1600 | 900 |  | No Backboards required. (Services required to be installed on benches with 1 mains lead with commado plug this will be hard wired into the RCD to protect the mains socket outlets. Plus data cables) | Floor | 4x Standard 240v switched sockets. | 8 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 2x Double Data Points
Mobile Single Bench 7 | NEW | Vender to Supply Mobile Benches (Services required See Note) | Vendor Install |  | 1530 | 1600 | 900 |  | No Backboards required. (Services required to be installed on benches with 1 mains lead with commado plug this will be hard wired into the RCD to protect the mains socket outlets. Plus data cables) | Floor | 4x Standard 240v switched sockets. | 8 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 2x Double Data Points
Double Fixed Bench 8 | NEW | Vendor to Supply Fixed Benches. Install supplies at Low level above benches in 3 way Dado trunking. | Vendor Install |  | 1530 | 4032 | 900 |  | Back boards to back and both ends | Floor | 4x Standard 240v switched sockets. | 8 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | 3x Double Data Points
Double Fixed Bench 9 | NEW | Vendor to Supply Fixed Benches. Install supplies at Low level above benches in 3 way Dado trunking. | Vendor Install |  | 1530 | 3193 | 900 |  | Back boards to back and both ends | Floor | 4x Standard 240v switched sockets. | 2 UNDER BENCH 2 ABOVE BENCH | N/A | N/A | N/A | YES COLD MAINS | N/A | N/A | N/A | N/A | N/A | N/A | YES | N/A |  | 2x Double Data Points
Double Fixed Bench 10 | NEW | Vendor to Supply Fixed Benches. Install supplies at Low level above benches in 3 way Dado trunking. | Vendor Install |  | 800 | 8434 | 900 |  | Back boards to back and both ends | Floor | 8x Standard 240v switched sockets. | 16 | N/A | YES TYPE 2 |  | YES 2x COLD MAINS 1x HOT SUPPLY | N/A | N/A | N/A | N/A | N/A | N/A | YES | N/A |  | 4x Double Data Points
Mobile Single Bench 11 | NEW | Vender to Supply Mobile Benches (No services required) | Vendor Install | Benches to have clear opening at frount to aalow for under bench cupboards. | 800 | 1600 | 900 |  | No Backboards required. No services required on these benches. |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Mobile Single Bench 12 | NEW | Vender to Supply Mobile Benches (No services required) | Vendor Install | Benches to have clear opening at frount to aalow for under bench cupboards. | 800 | 1600 | 900 |  | No Backboards required. No services required on these benches. |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Mobile Single Bench 13 | NEW | Vender to Supply Mobile Benches (No services required) | Vendor Install | Benches to have clear opening at frount to aalow for under bench cupboards. | 800 | 1600 | 900 |  | No Backboards required. No services required on these benches. |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Mobile Single Bench 14 | NEW | Vender to Supply Mobile Benches (No services required) | Vendor Install | Benches to have clear opening at frount to aalow for under bench cupboards. | 800 | 1600 | 900 |  | No Backboards required. No services required on these benches. |  | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Shelving | NEW | Vender to Supply & Install | Vendor Install |  | 2250 | 600 | 1500 |  | Back boards to back and both ends on all shelves. | Floor | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
Ladder Rack 1 | NEW | Vender to Supply & Install at low level with Services. | Vendor Install |  | 575 | 40 | Securly fixed to Floor & at Height |  | General | FLOOR TO ROOF | 8x Standard 240v switched sockets. Ref: Swing and Formax robots. | 16 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 3x Double Data Points
Ladder Rack 2 | NEW | Vender to Supply & Install with Services. | Vendor Install |  | 2050 | 40 | Securly fixed to Floor & at Height |  | General | FLOOR TO ROOF | 8x Standard double sockets. 2x 32A Single Phase 2x 32A 3~Phase N+E | Refer to power supply notes | YES TYPE 2 (1 EACH SIDE) | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 4x Double Data Points
Ladder Rack 3 | NEW | Vender to Supply & Install with Services. | Vendor Install |  | 2000 | 40 | Securly fixed to Floor & at Height |  | Supplies to be fitted equally on both sides of the ladder Rack | FLOOR TO ROOF | 8x Standard double sockets. 2x 32A Single Phase 2x 32A 3~Phase N+E | Refer to power supply notes | YES TYPE 2 (1 EACH SIDE) | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | 4x Double Data Points
Ladder Rack 4 | NEW | Vender to Supply & Install with Services. | Future |  | 3400 | 40 | Securly fixed to Floor & at Height |  | Supplies to be fitted at High level above for future installation. | FLOOR TO ROOF | 8x Standard double sockets. 2x 32A Single Phase 2x 3 Phase N+E | Refer to power supply notes | YES TYPE 2 | YES TYPE 2 | N/A | N/A | SOLVENT | 200 | 0.5 | N/A | N/A | N/A | N/A | N/A | N/A | 4x Double Data Points
Ladder Rack 5 | NEW | Vender to Supply & Install with Services. | Future |  | 2575 | 40 | Securly fixed to Floor & at Height |  | Supplies to be fitted at High level above for future installation. | FLOOR TO ROOF | 8x Standard double sockets. 2x 32A Single Phase 2x 3 Phase N+E | Refer to power supply notes | YES TYPE 2 | YES TYPE 2 | N/A | N/A | SOLVENT | 200 | 0.5 | N/A | N/A | N/A | N/A | N/A | N/A | 4x Double Data Points
2-off Emergency Showers | NEW | Vender to Supply & Install with Services. | Vendor Install |  | 1000 | 1000 |  |  |  | Floor | N/A | N/A | N/A | N/A | N/A | Hot / Cold Mains | N/A | N/A | N/A |  | N/A |  | Yes Standard Drain System | N/A | N/A | N/A
Wash Basin | NEW | Vender to Supply & Install with Services. | Vendor Install |  | 455 | 355 |  |  |  | Floor | N/A | N/A | N/A | N/A | N/A | Hot / Cold Mains | N/A | N/A | N/A |  | N/A |  | Yes Standard Drain System | N/A | N/A | N/A
Lab Sink 1 (R/H Drainer) | NEW | Vender to Supply & Install with Services. | Vendor Install |  | 1200 | 600 |  |  |  | Bench | N/A | N/A | N/A | N/A | N/A | Hot / Cold Mains | N/A | N/A | N/A |  | N/A |  | Yes Standard Drain System | N/A | N/A | N/A
Lab Sink 2 (R/H Drainer) | NEW | Vender to Supply & Install with Services. | Vendor Install |  | 1200 | 600 |  |  |  | Bench | N/A | N/A | N/A | N/A | N/A | Hot / Cold Mains | N/A | N/A | N/A |  | N/A |  | Yes Standard Drain System | N/A | N/A | N/A
2x Cordless Phone | NEW | New CPI to Purchase | New CPI to Purchase | [6 empty] | Bench | 1x Standard 240v switched sockets. | 2 | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | N/A |  | N/A | N/A | N/A | 1x Double Data Points
2x Full height double general storage cupboards | NEW | Vender to Supply | Vender to Supply |  | 1200 | 600 | 2000 |  |  | Floor | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | N/A |  | N/A | N/A | N/A | N/A
1x Full height double Flamables storage cupboards (With Bunded Shelf's) | NEW | Vender to Supply | Vender to Supply |  | 1200 | 600 | 2000 |  |  | Floor | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | N/A |  | N/A | N/A | N/A | N/A
16 x Under Bench Lab Storage / Single cupboard with lockable castors | NEW | Vender to Supply | Vender to Supply |  | 505 | 510 | 810 |  |  | Floor | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | N/A |  | N/A | N/A | N/A | N/A
24x Under Bench Lab Storage Double cupboard 2 doors with lockable castors | NEW | Vender to Supply | Vender to Supply |  | 1005 | 510 | 810 |  |  | Floor | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | N/A |  | N/A | N/A | N/A | N/A
4x Under Bench Lab Storage Double cupboard 4 draws with lockable castors | NEW | Vender to Supply | Vender to Supply |  | 1005 | 510 | 810 |  |  | Floor | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | N/A |  | N/A | N/A | N/A | N/A
2x Under sink Cupboards | NEW | Vender to Supply | Vender to Supply |  | 1200 | 600 |  |  |  | Floor | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A |  | N/A |  | N/A | N/A | N/A | N/A
0
Note on T7: David Bowron: Supplied from Vendor. M3 per/hr
Note on T30: David Bowron: waiting for information from Lynn
Note on A44: David Bowron: Lynn to advise