## Drop Codes (hidden)
RoomClassification | IA | QualificationTest
Not Classified | Direct Impact | Yes
Grade A | Indirect Impact | No
Grade B | No Impact | n/a
Grade C
Grade D

## Dispense System
User Requirement Specification |  |  | CPI Internal Tender Reference & Title:- | 1085 | Supply, Installation and Commissioning of an Optical Measuring Microscope
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 40%
USER REQUIREMENT |  |  |  |  | Tenderer Response
Requirement codes |  |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory |  |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  | Requirement |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
| System Features | Working area 400mm x 400mm |  | D
|  | Working area 350mm x 300mm |  | M
|  | Measuring accuracy +/- 10um |  | M
|  | Video & image capture and storage |  | M
|  | High resolution CCD video camera |  | M
|  | 4 trurret magnifcation array |  | D
|  | Magnification requirement x10 to x100 |  | M
|  | Magnification x200 to x500 |  | O
|  | Surface illumination (adjustable intensity) with colour lenses |  | M
|  | Rear illumination with adjustable intensity |  | D
|  | Moveable stage with holder for flexible films |  | M
|  | Automated moveable stage |  | D
|  | Ability to measure flexible substrates (opaque and transparent) with metallised patterns (copper, silver, aluminium) |  | M
|  | Ability to measure SMD 0201 size (0.6mm x 0.3mm) |  | M
|  | Ability to measure bare die silicon devices (0.3mm x 0.3mm) |  | M
Installation & Utilities
| Installation | Installation must be performed by qualifed engineers at CPI Newton Aycliffe facility |  | M
| Site requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment |  | M
| Electrical Power & utilities | The system should operate from a standard UK mains supply, other utilities to be listed |  | M
Software and Connectivity
| Connectivity | Computer Interface (e.g. ethernet, USB etc) |  | D
| Measurement | Measurement software with on screen capture and image & video storage |  | M
Regulatory, Support & Warranties
| CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided |  | M
| Training | Full training should be provided along with access to technical support |  | D
| Warranties | Warranty period should cover at least 18 months from installation |  | M
| Extended warranty | Extended warranty of up to 3 years |  | D
| Service support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. |  | D
| Service support | Details of preventative maintenance programmes required (if any) above Warranty requirements |  | D
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com