# Thurrock Council Specification

## INSTALLATION, SERVICING AND BREAKDOWN REPAIRS

## OF DISABLED ADAPTATION LIFTING AIDS

1.	INTRODUCTION AND BACKGROUND 	2

2.	AIM OF THE CONTRACT	2

3.	PRELIMINARIES 	2

3.1	Health and Safety	2

3.2	CDM Requirements	3

3.3	Customer Care	3

3.4	Working in Occupied Homes	4

3.5	Equality and Diversity	5

3.6	Working with the Council	5

3.7	Complaints	6

4.	REQUIREMENTS	6

4.1	Overview	6

4.2	Repairs, Maintenance and Installation overview	6

4.3	Asset Register	7

4.4	Servicing and Maintenance	8

4.5	Repairs Identified at point of Service/Inspection	12

4.6	Breakdown Repairs	13

4.7	Removal and Storage of Equipment for Re-use.	13

4.8	Supply and Installation of Lifting Equipment	13

4.9	Record Management	13

4.10	Service Visit Access Arrangements	14

4.11	Programme of Works for Servicing	14

5.	PERFORMANCE MANAGEMENT	14

5.1	Performance Management	14

5.2	Key Performance Indicators	15

5.3	Data Collection	15

5.4	Contractor’s Quality Assurance	15

6.	GENERAL DATA PROTECTION (GDPR) 	15

APPENDIX 1 – Key Performance Indicators 	16

### 1.	INTRODUCTION AND BACKGROUND

1.1	Thurrock Council has a housing stock of over 9,000 properties as a mixture of dwellings including high rise, low rise and individual houses.  Where required, the Council commission’s contractors to carry out disabled adaptation works to Council properties in order to facilitate independence for residents.  In addition, the Council manages the provision of the Disabled Facilities Grant (DFG) which enables owner occupiers in the borough to secure funding to carry out disabled adaptations to their properties where this is necessary.

1.2	This contract requires a single, suitably qualified Contractor to deliver Disabled Adaptations Services for the Supply, Installation, Servicing and Breakdown Repairs of lifts (stairs, step, and through floor) and Ceiling Track Hoists

The Contractor shall note that the start date of the servicing, breakdown repairs and installation elements will be 1 December 2019.

The terms and conditions of this contract will be the JCT Repairs and Maintenance Contract 2016.

1.3	Primarily the work will be delivered in Thurrock Council properties; however, there will be some work required in owner occupied properties, where this is managed directly by the Council.  Where owner occupiers choose to manage projects directly, they may choose alternative Contractors.

1.4	The scope and volume of work is subject to demand and the Council makes no commitment to value during the Contract Term.

1.5	In responding to this opportunity, Contractors confirm that they can comply with the relevant statutes and regulations current at the time of the selection process, including as relevant:

- Gas Safety

- Health and Safety

- Control of Asbestos,

- COSHH

- Electricity at Work

- Manual Handling

- Personal Protective Equipment.

### 2.	AIM OF THE CONTRACT

This Contract shall deliver high quality Disabled Adaptations Services to properties in Thurrock.  Achievement of the aims will generally be measured by the Contractor’s performance to this specification as set out in the Key Performance Indicators in Appendix 1 and Customer Satisfaction Surveys.

### 3.	PRELIMINARIES

#### 3.1	Health and Safety

3.1.1	All Contractors are expected to fully comply with ALL health and safety legislation.  All new operatives shall be formally trained in the safe use of all tools and Contractors shall ensure that their operatives have the correct equipment including personal protective equipment.  Full training and assessment records of all operatives must be kept and be available to view upon request.

3.1.2 	The Contractor and all persons (including sub-Contractors) employed by them on the work shall comply fully with:

- Health and Safety at Work Act 1974 and;

- All appropriate enactment’s which are relevant statutory provisions under the above Act

- All other relevant safety requirements

- Appropriate Codes of Practice

- Health and Safety Executives Good Practice Guidance Notes.

3.1.3	The Contractor shall provide all necessary equipment and safe provision for power to the works.

3.1.4		Unless explicitly agreed by the Authorised Officer, the Contractor will not be allowed to use the existing dwellings and their services and amenities. This shall apply irrespective of any permission and/or arrangements given or made between the Contractor and the customer at the dwelling.

3.1.5	Under the Control of Asbestos Regulations2012, the Contractor is expected, where Asbestos Containing Materials (ACMs) are discovered, to stop works immediately and notify the Authorised Officer without delay so corrective action can be taken. Under no circumstances shall any ACM be removed or disturbed, the requirement on the Contractor shall be to provide appropriate feedback immediately.

3.1.6	The Contractor shall put in place sufficient insurance to cover all on site risks both for their own equipment and staff, and for the resident and their property.

#### 3.2	CDM Requirements

3.2.1	Where Contracts are notifiable under CDM the Client will advise on the arrangement prior to the commencement of works in line with legislation.

#### 3.3	Customer Care

3.3.1	Thurrock have an excellent reputation for customer service, and aim to maintain this by working in partnership with their Contractors.

3.3.2	The Contractor shall follow best practice principles in their contact with residents regarding work to be completed in their property.  The following list provides examples of these principles, but is not exhaustive.

- All communication (e.g. letters) must be easy to understand and identify who is the contact (name, address, telephone number etc.)

- All workers must wear identity cards at all times

- The resident’s permission must be sought before the Contractor (or anyone acting for him) works/inspects etc. within the curtilage of the property or on the property itself.

- All workers will act as quietly as possible and take steps to minimise disruption to residents

- Work will take place in resident’s homes or within communal areas between the hours of 8am to 5pm.

3.3.3	Where the work also needs access to residents’ homes, the Contractor shall comply with the following additional requirements:

- Residents must be treated with respect and in a friendly and courteous manner.

- At all stages the resident must be consulted about the work.  For example, in making appointments to service gas appliances, making arrangements for the work to be done and deciding whether a satisfactory job has been done.

- The prescribed level of supervision by the Contractor must be guaranteed and, at an individual level, both client and Contractor representatives must establish good working relationships, understanding each other’s requirements.

- The Contractor shall update Service Users and client on the progress of breakdown repairs, for example, awaiting parts (time line to be included), awaiting approval of quotation etc.

3.3.4	In the interests of safeguarding, all employees will not have unsupervised access to children (under the age of 18) or vulnerable adults and it is important that the resident/householder maintains responsibility for supervision of persons on site during the works.

3.3.5	Contractors are not permitted to work in premises occupied by persons of under the age of 18 unless supervised by a member of the household of at least 18 years of age.

3.3.6	Where this supervision is not in place, the operatives must withdraw from site immediately until adequate supervision is reinstated,  record any delays, and report to the Contract Manager.

3.3.7	If the Contractor has any further concerns, including the behaviour of the resident/householder, they should report these immediately to the Employer’s representative.

#### 3.4	Working in Occupied Homes

3.4.1	Where the work is taking place in homes which are lived in whilst the work is carried out, the Contractor must ensure that the workforce engaged under this contract (including sub-Contractors where used) adopt a high standard of customer care at all times.

In addition to the customer service requirements set out in Section 3.3, all staff must adhere to the following rules as a minimum:

.

- The use of radios for site entertainment is not permitted

- Use of offensive or abusive language will not be tolerated

- All residents must be treated with proper respect and in particular the requirements of the elderly and those with special needs

- The Contractor shall prohibit staff from smoking or using any of the residents welfare facilities (WC, sink etc.) whilst attending any property; and

- All staff will be required to be properly and presentably dressed in appropriate work-wear to the satisfaction of the Authorised Officer.

3.4.2	The Contractor must treat the property of the customer with respect.  Specifically, it shall be brought to the Contractor’s notice that they shall be working around customer’s own white goods in confined areas and, therefore, must make all attempts to minimise disruption and damage.  The Contractor shall be liable for any damage he causes to the Resident’s own property by act or failure to act as appropriate.

3.4.3	The Contractor shall ensure that all residents have equal access to the service regardless of vulnerability or diversity. Thurrock will require formal acknowledgement that all operatives have undertaken recognised training with regard to Equality and Diversity

3.4.4 	Where necessary, dust sheets shall be used to minimise the impact of any works to the Resident’s property.

3.4.5	The Contractor, any workers, or sub-Contractors must not make any comment or pass opinions relating to housing maintenance/servicing decision or Council Policy.  Any queries about such a decision should be referred to the Contract Manager.

#### 3.5	Equality and Diversity

3.5.1	Thurrock Council are committed to equality of opportunity and wish to encourage Contractors to demonstrate a similar commitment. Contractors should therefore:

- Adopt their own equality and diversity policies and procedures;

- Take steps to ensure that they recruit people to reflect the ethnic mix of communities in which they are working;

- Abide by the statutory requirements of the Race Relations Act, the Sex Discrimination Act, the Age Discrimination Act and their respective Codes of Practice; and

- Where appropriate, seek advice from bodies such as the Council for Racial Equality (C.R.E) and the Equal Opportunities Commission (E.O.C).

3.5.2	Discriminatory behaviour or racial or sexual harassment towards its agents, employees, or Residents will not be tolerated. Any allegation of such behaviour will be investigated thoroughly by the Council and appropriate action will be taken if allegations are found to be sustained.

3.5.3	Thurrock Council Tenancy Agreements require Residents not to harass any neighbour, employee or agent of the Council. Contractors or their operatives and any Sub-Contractors who have suffered harassment should bring this to the attention of Thurrock Council who will fully investigate the allegations.

#### 3.6	Working with the Council

3.6.1	The Contractor shall deliver this service with the Council in a way that embodies the spirit of partnership.

3.6.2	The Council’s “Contract Manager” shall provide the overall client management element of the contract.  The “Authorised Officer” will provide the day-to-day liaison in terms of instruction, variation and data collection.

3.6.3	The Contractor, his operatives and sub-contractors must work under the ethos of “zero defects”, and each one of them must take responsibility for ensuring their works are free of defects prior to leaving the property.

#### 3.7	Complaints

3.7.1	The Contractor shall put in place a complaints system that is compatible with the Council’s Complaints Procedure, details of which are available at. https://www.thurrock.gov.uk/how-to-complain/general-complaints  Details of complaints and the remedial action where appropriate shall be shared with the Contract Manager at the performance meetings.

### 4.	SERVICING AND REPAIR REQUIREMENTS

#### 4.1	Overview

4.1.1	This contract is for the supply, installation, planned programme of preventative maintenance, and breakdown repairs of stairs, step, and through floor lifts and ceiling track hoists.

4.1.2	Contractors shall confirm that they can comply with the relevant Statutory Instruments and Regulations, current at the date of submission of tenders and in particular the following.

- Lifting Operations and Lifting Equipment Regulations(LOLER) 1998

- 17th Edition of the I.E.E Wiring Regulations plus Amendments

- Provision and Use of Work Equipment Regulations(PUWER) 1998

- BSEN 81-40:2008. Safety Rules for the Construction and Installation of Lifts

- BSEN 10535. Hoists for the Transfer of Disabled Persons. Requirements and Test Methods

- BS 5900: 1999 Specification for Powered Domestic Lifts or European equivalent

- Machinery Directive 2006/42/EC

#### 4.2	Repairs, Maintenance and Installation

4.2.1	Repairs shall be carried out on a responsive basis and new installations shall be left safe and fully operational.

4.2.2	Maintenance of lifts (stairs, step and through floor) shall be on an annual basis and ceiling track hoists shall be bi-annual, made up of one visual check and a load test.

4.2.3	The Contractor shall be required to carry out servicing and breakdown repairs of lifts and ceiling track hoists for Council residents.

4.2.4	Lifts installed in owner occupier properties under the Disabled Facilities Grant (DFG) shall be serviced and repaired by the installation Contractor, if in warranty. If out of warranty, owner occupiers will have to make their own arrangements for servicing and breakdown repairs.

4.2.5	The repair cost shall be based on the Schedule of Rates for breakdown repairs set out in the Pricing Schedule.  Where a faulty component is not on the schedule of rates, repairs shall be based on labour and material(s) cost.

4.2.6	Invoices for such repairs shall be broken down into material and labour cost.

4.2.7	The Contractor shall employ, under this contract, only Engineers who are competent (approved by equipment manufacturers) and with a level of expertise which will be required in installation, servicing and breakdown repairs of lifts (stairs, step and vertical) and ceiling track hoists of different design and manufacture.

#### 4.3	Asset Register

4.3.1	A comprehensive list of equipment, Service User names, addresses and phone numbers shall be provided to the successful Contractor.

4.3.2	At the point of this tender, the Asset Register is made up of the following:

| Equipment | Total Number | Number within Warranty |
| --- | --- | --- |
| Housing Properties |  |  |
| Stair Lifts | 335 | 44 |
| Ceiling Track Hoists | 108 | 8 |
| Through Floor Lifts | 16 | - |
| Step Lifts | 2 | - |
| Non Housing Properties |  |  |
| Stair Lifts | 1 | - |
| Step Lifts | 1 | - |
| Changing Benches | 7 | - |

4.3.3	The Contractor shall ensure that the Housing Lift Asset Register is verified during the first year of the contract period and any discrepancies made known, in writing, to the Authorised Officer.

4.3.4	Equipment within the optional Non Housing Properties (within Thurrock borough boundaries) are discretionary and the successful Contractor will be informed if these are to be included in the first or any year of the contract.

4.3.5	Equipment types within Thurrock include:

| Stair Lifts | Through Floor Lifts | Step Lifts | Ceiling Track Hoists | Changing Benches |
| --- | --- | --- | --- | --- |
| Minivator/Handicare 930,935, 950,950+, 950L, 1000 and 2000 Stannah 225, 240, 250, 260, 300, 400, 420, 420L and 420SL Prism Medical Simplicity, Horizon and Platinum Plus also: - Ambassador - Nimbus - Cumbria - Van Gogh - Bison Bede | Wessex VM Series Terry Pollock | Wessex LR800 and KL720 Pollock | Nopac 601- 603 Chiltern Sunrise Voyager Guildman BHMV series Prism TXM4 and TX270 | Chiltern |

4.3.6 	All Equipment under warranty with other Contractors will be added onto this contract at the end of the equipment’s warranty period.

4.3.7	The Servicing cost shall be same as that on Contractors pricing schedule.

4.3.8	Break down repairs cost shall be same as Contractors  Schedule of Rates or labour and material charges.

#### 4.4	Servicing and Maintenance

The Contractor shall carry out servicing of the equipment within normal working hours in accordance with Statutory Guidance and this Specification as generally described below and in accordance with the Manufacturers’ Requirements where these supersede.

##### 4.4.1	Stair Lifts – Straight or Curved

- Check motor/gear box for wear, overheating and leakage. Check drive pinion or chain sprocket for wear and correct meshing.

  - Check operation of motor brake and over speed governor.

  - Check condition of rollers.

  - Check all welds including roller brackets, main shaft bearing block and pivot/swivel on seat and footrest.

  - Check condition of rail, clean and lubricate track if applicable.

  - Check operation of manual/power hinged rail if applicable.

  - Check operation of sliding track rail if applicable.

  - Check all fixings are tight and correctly torqued including the chair, rail, rail hinge (where applicable) and limit stops.

- Check all clips, split pins etc. are correctly fitted.

- On curved rail lifts, check condition of trailing cable, correct reeling, drum and pulleys are secure and lubricated.

- Check correct operation of all switches, chair and landing controls, all safety gear contacts, time delays, and safety edge contacts. Check condition of wiring.

- Check operation of RCD where fitted. Check spur where fitted.

- Check all earth terminals on rail, carriage and trailing cable. Check earth continuity on relevant points.

- Check start and run current for up and down travel on lift. On DC battery lifts, check for correct charging and charger L.E.D’s are illuminated. Check power chain cable if fitted.

- Check condition and operation of hinged platform if fitted. Check condition of footrest rubber and chair upholstery.

- Check bottom and top call station controls.

- Check operation of seat belt/harness.

- Check operation of footplate.

- Check joints in track.

- Leave a laminated notice, reminding service user or carer to leave their equipment on charge at all times.

- Clean charge strip/points.

- Check operation of safety edges during upward and downward travels.

##### 4.4.2	Ceiling Track Hoists

##### Check spreader bolt.

- Check operation of spreader bar.

- Check spreader bar pads.

- Check track fixings.

- Check battery and any leads are fitted correctly.

- Check track end stops

- Check operation of traverse wheel drive.

- Check trolley wheels.

- Check over speed protection.

- Check up/down limit switches.

- Check wiring and fuses.

- Check manual traverse.

- Check batteries main unit.

- Check emergency lowering (in the event of power cut).

- Check back up unit.

- Check hand control.

- Check condition of rollers.

- Check charger.

- Check isolation switch.

- Check emergency stop.

- Check audible overload warning and cut out, where applicable.

- Check labelling/date of visit.

- Check transformer.

- Check sling is clean and undamaged.

- Check label on sling is legible with SWL clearly displayed.

- Check unique identifier on sling.

- Carry out a load test.

- Check transformer.

- Completely unwind strap/lifting tape from drum and check for damage, signs of wear or discolouration.

- Clean inside of track at every visit.

- Clean hoist with a damp cloth using warm water and disinfectant.

- Check LDR reading.

##### 4.4.3	Through Floor Lift

Lift to be serviced to BS 5900 and includes checks on the following:

- Safety interlocks checks.

- Hydraulic leakage checks (e.g. hoses, fittings etc.)

- Hydraulic pressure and safety valve checks.

- Screws and nuts (as applicable).

- Suspension member checks.

- Safety circuits, fuses and other electrical equipment.

- Lubrication checks.

- Check wall mounted control stations at ground and floor levels.

- Check car lighting.

- Check car sensitive edges/surfaces.

- Check grab handles where fitted.

- Check seat belt where fitted.

- Check emergency lowering controls in lift car.

- Check emergency lowering by pump unit control

- Check safety gear and over speed governor (arrangements for preventing fall of the car and the counter weight).

    - Check suspension ropes and their attachments, if applicable.

- Check over running devices.

- Check car fittings, car and counterweight guides, if applicable.

- Check interlocks on car doors.

- Check car doors.

  - Check other door fastenings.

##### 4.4.3	Step Lift

- Check sensitive edge assembly.

- Check actuator assembly.

- Check Ramp lock assembly.

- Check pump box.

- Check controls.

- Check hydraulic ram assembly.

- Check bottom stop switch.

- Check top stop switch.

- Check Ramp Assembly

- Check left hand side assembly.

- Check right hand side assembly.

- Check scissor frame where applicable.

- Check general condition of lift.

#### 4.5	Repairs Identified at point of Service/Inspection

4.5.1	Where equipment is found to be at all defective during a service visit, faulty components shall be replaced immediately, but cost of repair must not exceed £150 (excluding VAT) and shall be reported to the Authorised Officer via telephone or email and the details shall be included on the service report.

4.5.2	Where a replacement part is not immediately available or possible, or the value of the proposed repair exceeds £150 (plus VAT), the Contractor shall notify the Authorised Officer and arrangements made as soon as possible to repair.

4.5.3	Unless it can be immediately repaired, equipment found to be unsafe but still operational shall be disabled and a warning notice “Do Not Use. Unsafe Equipment” shall be attached. The Contractor shall notify the Authorised Officer the same day.

4.5.4	The Contractor shall ensure that all engineers maintain a stock of suitable spares to cover minor repairs valued up to £150 excluding VAT to enable the systems to be repaired and left operational without further need to secure parts.

4.6	Breakdown Repairs

4.6.1	Breakdown Repairs will be ordered by the Authorised Officer using the following timescales.

| Call Out Period | Response Time | Access Arrangements |
| --- | --- | --- |
| Normal Working Hours 8am to 5pm Monday to Friday | 2 Hours of instruction | As this covers visits to residents’ homes, any specific details of resident vulnerability will be communicated to the Contractor at the point of call out. |
| Out of Hours 5pm to 8am Monday to Friday, and all weekends and public holidays | 4 Hours of instruction |  |

4.6.3	If in the opinion of the Council the breakdown is due to faulty materials, workmanship or design of components provided by the Contractor, the cost of such shall be met by the Contractor.

#### 4.7	Removal and Storage of Equipment for Re-use

4.7.1	The Council may require lifting equipment to be removed from properties, and stored by the Contractor for re-use at a later point.

4.7.2	The Contractor must allow for storage of all reusable equipment with various spare parts for the duration of the contract. An accurate account of stored equipment must be supplied to the Authorised Officer at the end of each month.

#### 4.8	Supply and Installation of Lifting Equipment

4.8.1	This Contract Lot includes the supply and installation of stair lifts, ceiling track hoists, step and through floor lifts.  The precise requirement for the installation will be agreed between the Authorised Officer and the Contractor.  The Contractor shall allow for a site visit and design of installation prior to agreement of the order.

4.8.2	All equipment shall be installed and commissioned in accordance with manufacturers’ requirements and subject to warranty.  The Contractor shall ensure that he is able to supply a wide range of equipment in order that the most suitable solution is provided for residents.

4.8.3	In some properties building works will be required to complete the installation.  The Contractor shall ensure they have the capability to carry out these works to the high standards required by the Council and residents.  The final quotation of any installation shall include the cost of such works.

#### 4.9	Record Management

4.9.1	Provision of accurate up-to-date records is a key element of contract performance and the Council’s ability to ensure that maintenance has been carried out within timescales and to best industry practice, and, where appropriate, in compliance with any relevant statutory legislation.

4.9.2	The Contractor shall provide and maintain the following records:

4.9.2.1	Service report sheets shall be designed by the appointed contractor and submitted for approval by the Authorised Officer.

4.9.2.2	Service report must confirm that equipment is safe to use/operate

4.9.2.3	At the end of each month, using electronic mail, contractor shall be required to provide the Authorised Officer with a list of equipment serviced and repaired during the month along with the reports. Forwarded on or before the 10th of the following month. For example, service and break down reports for June to be forwarded on or before 10th July.

4.9.2.4	Repairs cost for each month shall be provided by contractor using an Excel spread sheet issued by the Authorised Officer. Where addresses are missing from the spread sheet, contractor must update the spread sheet accordingly.

4.9.2.5	Service and breakdown reports shall be forwarded electronically.

4.9.2.6	Upon completion of a service, contractor will place a sticker on the equipment. Sticker will have date of service, next service date and service engineer’s name.

4.10      Service Visit Access Arrangements

4.10.1	All service visits must be by prior arrangement with the service user with reasonable notice given and contractor must be flexible when making arrangements for a service visit as service users are vulnerable and may have other appointments (hospital, dental etc.).

4.10.2	Where contractor is having access problems using electronic mail, these should be brought to the attention of the contract administrator with time and date of calls and correspondence to the service user.

4.11     Programme of Works for Servicing

The contractor will be required to provide a programme of works for servicing of equipment based on the asset register provided. This will be further discussed at the pre-start meeting.

### 5.	PERFORMANCE MANAGEMENT

#### 5.1	Performance Management

5.1.1	The Council’s Authorised Officer will be responsible for Performance Management of the Contract.  Management will take place through analysis of data, consideration of performance against the KPIs and regular meetings with the Contractor.

5.1.2	Contract Management meetings shall take place at regular intervals between the Council and the Contractor, according to need and usage.

5.1.3	Also regular progress meetings will be held at the Employer’s representative offices. The contractor is to arrange for a representative to attend.

#### 5.2	Key Performance Indicators

5.2.1	The Contractor shall provide data to evidence their performance against the Key Performance Indicators set out in Appendix 1.  These KPIs are set for the first year of the contract.  Detailed report to be submitted to the Contract Manager and should be to the schedule as set out in the “Reporting Frequency” column.

5.2.2	The Council will measure both the performance and quality against the relevant Key Performance Indicators on an annual basis and agree with the Contractor any revisions of these in advance of the next contract year.

#### 5.3	Data Collection

Additionally, the Contractor shall provide data as requested by the Contract Manager which will evidence further performance of service delivery.  This data may not initially have a target figure, but may form part of the discussions and amendments to the list of Key Performance Indicators.

#### 5.4	Contractor’s Quality Assurance

The Contractor shall put in place a robust quality management system that they will use for internal monitoring to ensure that the level of service delivered is as required by the Council.  Details of the proposed quality Assurance System shall be provided within the tender submission.

6.	GENERAL DATA PROTECTION REGULATIONS (GDPR)

The Service Provider will comply with all aspects of the Data Protection Act 1998 (DPA) and the following Council policies (attached):

- Data Protection Policy

- Information Security Policy

- Records Management Policy

The Service Provider may act as a ‘Data Processor’ on behalf of the Council. Whilst Service Providers currently have certain responsibilities under the Data Protection Act, the new GDPR will bring increased responsibility and accountability for data processors.

The winning Bidder of this contract will be expected to complete a “General Data Protection Form – Light Touch Questions” and “Data Processor Agreement” prior to commencement of the contract as the data under this contract will be Service User’s name, address and contact telephone number only.

APPENDIX 1 - KEY PERFORMANCE INDICATORS – Installation, Repair & Servicing of Disabled Adaptations for Lifting Aids

| Objective | Output | Measurement | Target | Reporting Frequency |
| --- | --- | --- | --- | --- |
| SERVICE DELIVERY | a) All Equipment Servicing to be up to date | 1a) Percentage of Servicing reports received within 2 weeks of the end of the month in which the service was scheduled divided by the percentage of reports received within 2 weeks 2a) Number of services completed per month divided by the number of required services to be completed per month. | 100% 100% | Quarterly Monthly |
|  | b) Breakdown repair resolved within “Normal hours” (within 2 hours of request) | b) Number of “Normal hours” repairs resolved divided by the number of “normal hours” repairs required | 100% | Quarterly |
|  | c) Breakdown Repair resolved “Out of hours” (within 4 calendar days of request | c) Number of “Out of hours” repairs resolved divided by the number of “Out of hours” repairs required | 95% | Quarterly |
|  | d) Repairs value up to £150 completed at first visit | d) Number of Repairs up to £150 completed by first visit divided by the number of up to £150 repairs required | 100% | Quarterly |
|  | e) Repair value >£150 completed within 1 week of Authorisation | e) Number of repairs >£150 completed within 1 week divided by the number of repairs >£150 required | 95% | Quarterly |
|  | f) Up to date Asset Register | f) Number of up to date Asset registers provided divided by the number of up to date Asset Registers required | 100% | Annually |
|  | g) Accurate Account of Stored Equipment | g) number of up to date storage records provided within 1 week divided by the number of up to date records required within 1 week | 100% | Monthly |
| Complaints | To minimise complaints received from the public | Number of complaints received per quarter | <3 | Quarterly |
| Value for Money | Expenditure compliant with Contract price | Percentage of non-compliant invoices divided by the number of Invoices compliant with agreed prices | 100% | Monthly |
| Performance Review | Performance Review Meetings | Percentage of Attendance at the Review Meetings divided by the number of Review Meetings required | 100% | Quarterly |
| Social Value | Achievement of Agreed Social Value Outcomes | TBC |  | Annually |