Specification Requirements

Supply, delivery and installation of a Micro-Macro hardness testing machine

For: Engineering Materials, HSE-Science Division; Major Hazards Capability Group - V04/02/21

| Area | Requirements | Explanatory notes |
| --- | --- | --- |
| Basic information | - Micro-Macro hardness testing machine for the hardness testing of metals - Must be capable of carrying out Vickers Hardness testing and Vickers micro-hardness testing of prepared metal samples - Must be Bench mountable - Must comprise rotating turret capable of taking at least one hardness indenter and at least two magnifying objectives |  |
| Hardness types | - Essential requirement is ability to measure Vickers hardness and Vickers micro-hardness - Desirable is the ability to also measure Brinell and Knoop hardness |  |
| Indenters | - Minimum requirement is Vickers hardness indenter - Desirable that the machine can also be used for Brinell and Knoop hardness with additional indenters. - Please state how many different indenters the turret can be loaded with at any one time | In the case that different indenters can be used, please state the cost of a Brinell indenter and a Knoop indenter separately. |
| Load range | - Essential load range is 10 g to 30 kg - Desirable minimum load is 1 g - Desirable maximum load is 60 kg |  |
| Objectives | - Machine must be provided with both x10 and x50 magnification objectives - These must be located in a turret system enabling easy swapping between magnifications - The machine must have the facility for fitting of other additional higher magnification objectives | In the case that different objectives can be fitted, please state the cost of a x100 objective separately. |
| Stage size and control | - Stage must have automatic control - Option for manual stage control is desirable - Minimum stage size of 200 mm x 200 mm - Larger stage size is desirable - Controllable in X and Y directions |  |
| Sample levelling and retention | - Any jigs or holders required to ensure sufficient restraint and levelling of the sample must be included in the total price |  |
| Indentation positioning | - Equipment must have options for manual and automatic positioning of indentations - Ability to programme indentations at set locations along a line in the X and Y directions | Ability to programme a 2D grid of indentations at set spacings in both directions is desirable |
| Hardness value calculation | - Automatic measurement of indentations and calculation of hardness measurements is required |  |
| Photographic Imaging | - Ability to capture images of specimen surface with indentation locations is required - Ability to capture and transfer such photographic files in standard formats (JPEG etc) is required |  |
| Dimensions and mass | - Please state width and length of unit - Please state foot print of machine - Please state mass of machine with all attachments | Unit must fit on bench top of width 850 mm |
| Installation and setup | - Installation, set-up and training (within current limits of access/distancing restrictions) must be included as part of price | Installation location is Buxton, Derbyshire |
| Software | - Please include all necessary standard software and control systems (PC, touch screen etc) in the quoted price - Please state whether software maintenance and updates are FOC. If not please state approximate annual cost of software maintenance fee. |  |
| Warranty | - Please state warranty period |  |
| Lead time | - Please state best lead time from receipt of order |  |

Bids will be evaluated and awarded on a Most Economically Advantageous Tender (MEAT) basis, including the ability to deliver before 31 March 2021.