Particle Sizer Specifications

Durham University

Department of Chemistry

Equipment required: Particle Sizer

Category: Laboratory equipment

Description: a light-scattering instrument for characterizing particles in liquid dispersions, which measures: particle size; zeta potential; electrophoretic mobility; molecular mass; transmittance.

Precise Specification Required:

Particle Sizer

- Three different measurement angles (forward-, side- or back-scattering)

- Measures transmittance continually

- Allows automatic adjustment of parameters (including at least: focus, attenuation, measuring angle, measurement duration)

- Zeta potential and mobility measurements

- DLS size measurement range from 0.3 nm to at least 10μm particle diameter

- Zeta potential range > +/- 600mV with size range from 3.8 nm to 100μm particle diameter

- Molecular mass range 980Da to 20000kDa

- Ability to measure with organic solvent cell

- Temperature control

- Windows 10 compatibility

- 3 year warranty with at least 2 planned service maintenance visits

- Installation and Training in how to use the equipment

Operational Software for the Particle Sizer

- Ability to see all important information on a single display, including input parameters, results, and final calculated values

- Ability to access measurement series, databases for solvents, materials and SOPs

- Ability to compare and analyse results from separate experiments, and to present these in customisable reports

- Ability to export results to an Excel file for further processing and analysis

- License for unlimited installations

Attachments

- Organic Solvent Cell: a cell allowing zeta-potential and DLS measurements in organic and aqueous solvents

Desirable/Optional Specifications:

- Provides direct measure of turbidity

- Dosing/flow module (please quote price separately)