## Drop Codes (hidden)
RoomClassification | IA | QualificationTest
Not Classified | Direct Impact | Yes
Grade A | Indirect Impact | No
Grade B | No Impact | n/a
Grade C
Grade D

## URS Document
User Requirement Specification |  |  | CPI Internal Tender Reference & Title:-
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 40%
USER REQUIREMENT |  |  |  |  | Tenderer Response
Requirement codes |  |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as Desirable) |  |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  | Requirement |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
|  | Standard USP1 and II compatible apparatus with friction free stirring |  | M
|  | Ability to measure dissolution of non-oral formulation types |  | D
|  | Able to use a range of smaller volume vessels (e.g. 50 mL - 1000 ml) |  | M
|  | Fits standard apparatus used by industry (i.e. Distek) or is supplied with a compatible apparatus |  | M
|  | Can use online measurement techniques to avoid necessity for offline HPLC and can achieve rapid measurements (e.g. fiber optics) |  | M
|  | Autosampling capabilities for off-line UV or HPLC analysis. |  | D
|  | Ability to translate dissolution methods from partner companies |  | D
|  | Small benchtop footprint with 6 to 8 dissolution baths |  | M
|  | Easy cleaning for potent and non-potent drug molecules |  | M
|  | Computer control from IT supplied by the vendor |  | M
|  | Upgrade to 12 channel system possible |  | D
|  | Option to purchase a validation tool to check bath fill level and the stirrer speed |  | O
|  | Flow through cell dissolution apparatus |  | D
|  | Able to visualise whole formulations via UV imaging |  | M
Installation & Utilities
| Installation | Installation must be performed by qualifed engineers at CPI NetPark facility |  | M
| Site requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment |  | M
| Electrical Power | The system should operate from a standard UK mains supply (240Vac, 50Hz, single phase) |  | M
| Other Utilities | The system will not require a 3 phase supply, cooling water, gases or cryogenic gases |  | D
Software and Connectivity
| Connectivity | Computer Interface (e.g. ethernet, USB etc) |  | D
Regulatory, Support & Warranties
| CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided |  | M
| Training | Full training should be provided along with access to technical support |  | D
| Warranties | Warranty period should cover at least 12 months from installation |  | M
| Extended warranty | Extended warranty of up to 5 years |  | D
| Service support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. |  | D
| Service support | Details of preventative maintenance programmes required (if any) above Warranty requirements |  | D
OFFICE USE ONLY - This section below this line is for CPI Evaluation and Internal USE only
| Notes for Evaluation Team: |  | Name of Bidder to be inserted here | 0 | 0
Please assess all of the answers given, with particular attention to the Mandatory answers provided, taking into account, the Desirable and Optional answers too, however, you do not need to score each individual answer, but based on the scoring range 0-5, just give one singluar score (default set at "5" in the red and yellow box below, and as we have to provide feedback to bidders, you must also provide comments on which section(s) above where you felt that they did not meet the requirements.
| Evaluation Panel |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com