Schedule A

### THE HEALTH AND SAFETY EXECUTIVE

STATEMENT OF REQUIREMENTS FOR THE PROVISION OF SUPPLY OF A MULTIFUCTIONAL MICROPLATE READER T07 1516

1	THE HEALTH AND SAFETY EXECUTIVE (HSE)

1.1 	The Health and Safety Executive (HSE) is a Crown non-departmental public body with specific statutory functions in relation to health and safety. It is appointed by the Secretary of State for Work and Pensions and employs around 3500 staff including policy advisers, inspectors, technologists and scientific and medical advisers. HSE’s job is to prevent people being killed, injured or made ill by work.

  - HSE consists of a governing Board comprising of a Chair and nine non-executive members, all of whom are appointed by the Secretary of State for Work and Pensions after consultation with organisations representing employers, employees, local authorities and others, as appropriate.

1.3 	HSE comprises various directorates and groups and is led by a senior management team.  The Health and Safety Laboratory (HSL), Chemicals Regulation Directorate (CRD) and Office for Nuclear Regulation (ONR) are also part of HSE.  HSE works from over 30 locations throughout Great Britain.

2.	Health & Safety Laboratory

2.1 		The Health and Safety Laboratory (HSL) is Britain's leading industrial health and safety facility. Operating as an agency of the Health and Safety Executive (HSE), HSL supports HSE’s mission to protect people's health and safety by ensuring risks in the changing workplace are properly controlled.

2.2	HSL are inviting tenders for the supply of a multifunctional plate reader with colorimetric, fluorescence and luminescence mode equipped with an environmental atmosphere control chamber suited for toxicological tests on living cells. Part of our requirement  is that new the instrument provides a multichromator mode for the fluoresence measurements and Spectral scanning for the light absorbance.

The new instrument is replacing a 12 year old mutlifunctional platereader which can no longer be supported by the manufacturer.

3	BACKGROUND

3.1		HSL provides a range of analytical services to external customers which enable them to assess exposure to biohazards. These services are essential to underpin companies risk assessments for consumer products and for protecting the health of those working with these materials. Over the last ten years HSL has added bespoke toxicological assay services which customers use to determine whether substances in their manufacturing process are potentially hazardous to health.

The analytical immunoassay and enzyme monitoring services that HSL provide are dependent completely on the use of micro-titre plate readers to control chemical and biological reactions and to quantity analytes in large numbers of samples (high throughput). The existing Tecan multifunctional plate reader purchased 12 years ago requires more manual optimisation of assays adding at least 5% more time than should be required and should it break down is not supported for maintenance by the manufacturer.

4	OBJECTIVES

4.1. To underpin HSL analytical service for quantifying biohazards (allergens, toxins and micro-organisms)

5	SPECIFICATION REQUIRED

5.1.	The key requirements for this equipment are as follows

- A multi-functional microwell plate reader that can be used for fluorescence, light absorbance and luminescence measurements

- Capability to read reactions in micro-well formats up to 384 well plates

- To provide for top and bottom measurement modes

- Suitable for endpoint and kinetic measurements

- Providing for individual well spectral scanning in the fluorescent and light absorbance modes

- To provide for UV and visible light absorbance modes

- To provide as standard modes fluorescent intensity measurements, fluorescence polarisation, and time resolved fluorescence.

- To provide for luminescence modes as flash and glow

- To provide a spectral range from approximately 350 to 740 nm

- To provide for spectral scanning between 250 and 1000 nm

- Read height optimisation

- To enable multi spot readings within wells in order to assess heterogeneous reactions

- To allow the use of additional filters for operating in the fluorescence mode

- To provide temperature control from at least 4oC up to 45oC with options for higher temperature above 60oC

- To provide plate shaking modes (linear, orbital and user defined)

- To provide environmental control of atmospheric gases (O2 and CO2) suitable for aerobic culture of mammalian cells and anaerobic culture of micro-organisms.

Other desirable factors

- Training support for staff to enable them to start using the instrument

- Ongoing technical support

5.2 	Instrument

The supply of the instrument will include all components hardware and software and appropriate electrical connection etc. The instrument needs to be a compact bench top model with the following features:

- Micro-injector system and controls suited to time kinetic assay and acute toxicology assays

- Control software that provides for setting up and saving assay protocols and changing all of the measurement parameters. Optionally this should be compliant with FDA regulation to ensure the security of the assay protocols and settings.

- The environmental control unit (included temperature and atmospheric gas levels) needs to allow cultured cells to be maintained stably up to 48hours for the purppose of undertaking toxocity assays.

- Data analysis software that enable a variety of mathematical functions (averages, blank subtractions, coefficient of variations, standard deviations etc.,) and statistical calculations. The software should allow data visualisation (graphs, trend analysis, equations solving and line fitting) as well as for analysing kinetic analysis data (for example enzyme kinetic line fitting). It should provide options for exporting raw of formatted (tabulated data) to standards office and analytical software packages used for generating reports or presenting the results.

- The software license should allow for use of several computers to allow data transfer and analysis within HSL.

5.3	 Installation, Commissioning, Training

Tender to include installation commissioning and delivery to HSL Buxton Derbyshire and onsite training.

5.4	 Warranty and Documentation

Customary, warranty and / or extended warranty to be offered.

All user manuals and copies of software and manuals should be received by HSL with the system hardware.

On warranty HSL would like to achieve 24 month warranty period as standard for the instrument.

6.	Information Technology

HSL has a separate Facilities Management IT Provider whom are responsible for the whole network of the laboratory. In that regard, they have a clear responsibility for ‘impact testing’ new IT equipment onto HSL’s IT network. The laboratory’s FM IT Provider sets a minimum requirement on all new equipment.

It is therefore a requirement that all Tenderers should specify in the fullest detail possible, the makers name and specification etc of all IT equipment (hardware, software etc.) that they propose to supply including details of the proposed operating system. Details of any associated IT peripherals and software should also be supplied. It would greatly assist if separate itemised costs for all of the associated IT components (e.g. PC, LCD, and Printer) could be included within the tender response, should these be available.

In addition, please provide warranty information, stating whether it is standard warranty (return to base) or onsite warranty (next business day), together with the length of warranty that will be supplied with the proposed equipment. Please also provide details of exactly how these warranties shall be activated and/or requested, should any problems occur.

PC Minimum Specification:

Processor: Intel i5 (or equivalent)

Memory: 4GB

Hard disk: 500GB

Display: DVI capable of 1280 x 800 resolution

Operating System: Windows 7 Professional

Connectivity: Network adapter (Ethernet RJ-45)

The PC is considered as part of the 'system' we buy and should be specifically included in any warranty agreement offered by the tenderer. We would expect that any problems or repairs with the software and/or hardware that forms part of the 'system' should be resolved by the supplier.

- Software

Prior to placing any order, HSL requires assurance from the vendor that the proposed software contained within the order does not require administration rights to run and operate correctly post installation. This is because HSL is unable to run and operate software in MS Windows that needs to have administration rights before it will run and operate correctly post installation. Therefore we wish the vendor to:

- Confirm that the software does not require administration rights to run; or

- If you are unable to provide a version of software that does not run and operate correctly without “administration rights”, please advise what you can do to resolve this.

Note that in the event of an order being placed, and it being found that “administration rights” are required to run and operate the software, or, alternatively, that you do not notify HSL's Purchasing Procurement Unit that the software needs administration rights in order to run and operate correctly, HSL reserves the right to cancel the order and to place an order with another supplier and to seek a refund of any payments that have been made against the purchase order.