## Lot 1
Guidance for completion - When inserting additional rows in the requirements table below (Starting at row 16), please ensure that you drag the formulas from Column A down from the row above into the newly inserted row. - Do not delete the top and bottom rows of the requirements table (Starting Row 16 and finishing Row 48). Only insert rows between these points and do not add any lines after these points. - Hide row 5 if lots arent applicable
| User Requirement Specification | CPI Tender Reference & Title:- | 1924 | Framework Supply and Delivery of High Temperature Ovens, Aftersales & Maintenance
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column C, of which the total Quality element score of this sheet will equate to value in cell F6 | Scoring Criteria 0 = no response/answer not relevant to requirement 1 = limited response to requirements 2 = less than adequate, filfils some of the basic requirements 3 = adequate response, fulfils all basic requirements 4 = good response, fulfils all requirements to a high standard 5 = excellent response, fulfils requirements to an exceptionally high standard |  | The total percentage available which will contribute towards the Quality scoring element of this tender | 40%
USER REQUIREMENT |  |  |  | TENDERER RESPONSE
Requirement Codes |  |  |  | Guidance
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as desirable) |  |  | Tenderers must provide responses within the following columns; 1. Column F - please provide details of your equipment functionality in relation to CPI requirements specified in column E 2. Column G - please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response |  | Y
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  |  | N
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
# | Parameter | Requirement | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N) | N
System Performance
M1 | General oven design | Ovens designed to meet and comply with BS EN 61010-2-010:2003 | M |  |  | 1 | 0 | 0
M2 |  | Oven versions designed to meet and comply with BS EN 1539:2015 | M |  |  | 2 | 0 | 0
M3 |  | Over temperature trips installed as standard | M |  |  | 3 | 0 | 0
O1 |  | Digital controllers programmable to allow multistep heating profiles | O |  |  | 3 | 1 | 0
D1 |  | Forced convection system | D |  |  | 3 | 1 | 1
D2 |  | Option for cable entry ports | D |  |  | 3 | 1 | 2
D3 |  | Ovens able to take multiple shelves (at least 2) | D |  |  | 3 | 1 | 3
M4 |  | Able to supply wide variety of oven types and configurations in terms of operational temperatures and volume capacity. | M |  |  | 4 | 1 | 3
D4 | Other System Requirements | Volume capacity greater than 50 litres | D |  |  | 4 | 1 | 4
D5 |  | Maximum operational temperature 300oC | D |  |  | 4 | 1 | 5
D6 |  | Temperature stability +/- 0.5oC or better | D |  |  | 4 | 1 | 6
D7 |  | Temperature uniformity +/- 7oC (at 300oC) or better | D |  |  | 4 | 1 | 7
0 | Installation and Utilities |  |  |  |  | 4 | 1 | 7
M5 | Installation | Installation must be performed by qualifed engineers at CPI facility | M |  |  | 5 | 1 | 7
M6 | Site Requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment | M |  |  | 6 | 1 | 7
D8 |  | Ovens should be able to sit on standard laboratory bench, depth 600 mm | D |  |  | 6 | 1 | 8
M7 | Electrical Power | The system should operate from a standard UK mains supply (240Vac, 50Hz, single phase) | M |  |  | 7 | 1 | 8
M8 | Other Utilities | The system will not require a 3 phase supply, cooling water, or cryogenic gases | M |  |  | 8 | 1 | 8
D9 |  | Versions of the ovens should be available designed so that they can operate under inert atmospheres (e.g. Nitrogen gas) | D |  |  | 8 | 1 | 9
0 | Software and Connectivity |  |  |  |  | 8 | 1 | 9
D10 | Connectivity | Computer Interface (e.g. ethernet, USB etc) | D |  |  | 8 | 1 | 10
0 | Regulatory, Support and Warranties |  |  |  |  | 8 | 1 | 10
M9 | CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided | M |  |  | 9 | 1 | 10
M10 | Manuals | Instrument and controller manuals must be supplied | M |  |  | 10 | 1 | 10
O2 | Training | Where appropriate full training should be available along with access to technical support | O |  |  | 10 | 2 | 10
M11 | Warranties | Warranty period should cover at least 18 months from installation | M |  |  | 11 | 2 | 10
M12 | Extended Warranty | Extended warranty of up to 5 years | M |  |  | 12 | 2 | 10
M13 | Service Support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. | M |  |  | 13 | 2 | 10
M14 |  | Details of preventative maintenance programmes required (if any) above Warranty requirements | M |  |  | 14 | 2 | 10
[6 empty] | 14 | 2 | 10
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: CHANGE THIS IN ACCORDANCE WITH THE TENDER DOC

## Lot 2
Guidance for completion - When inserting additional rows in the requirements table below (Starting at row 16), please ensure that you drag the formulas from Column A down from the row above into the newly inserted row. - Do not delete the top and bottom rows of the requirements table (Starting Row 16 and finishing Row 48). Only insert rows between these points and do not add any lines after these points. - Hide row 5 if lots arent applicable
| User Requirement Specification | CPI Tender Reference & Title:- | 1924 | Framework Supply and Delivery of High Temperature Ovens, Aftersales & Maintenance
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column C, of which the total Quality element score of this sheet will equate to value in cell F6 | Scoring Criteria 0 = no response/answer not relevant to requirement 1 = limited response to requirements 2 = less than adequate, filfils some of the basic requirements 3 = adequate response, fulfils all basic requirements 4 = good response, fulfils all requirements to a high standard 5 = excellent response, fulfils requirements to an exceptionally high standard |  | The total percentage available which will contribute towards the Quality scoring element of this tender | 40%
USER REQUIREMENT |  |  |  | TENDERER RESPONSE
Requirement Codes |  |  |  | Guidance
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as desirable) |  |  | Tenderers must provide responses within the following columns; 1. Column F - please provide details of your equipment functionality in relation to CPI requirements specified in column E 2. Column G - please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response |  | Y
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  |  | N
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
# | Parameter | Requirement | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N) | N
System Performance
M1 | General oven design | Ovens designed to meet and comply with BS EN 61010-2-010:2003 | M |  |  | 1 | 0 | 0
M2 |  | Oven versions designed to meet and comply with BS EN 1539:2015 | M |  |  | 2 | 0 | 0
M3 |  | Over temperature trips installed as standard | M |  |  | 3 | 0 | 0
O1 |  | Digital controllers programmable to allow multistep heating profiles | O |  |  | 3 | 1 | 0
D1 |  | Forced convection system | D |  |  | 3 | 1 | 1
D2 |  | Option for cable entry ports | D |  |  | 3 | 1 | 2
D3 |  | Ovens able to take multiple shelves (at least 2) | D |  |  | 3 | 1 | 3
M4 |  | Able to supply wide variety of oven types and configurations in terms of operational temperatures and volume capacity. | M |  |  | 4 | 1 | 3
D4 | Other System Requirements | Volume capacity greater than 25 litres | D |  |  | 4 | 1 | 4
D5 |  | Maximum operational temperature 600oC | D |  |  | 4 | 1 | 5
D6 |  | Temperature stability +/- 0.5oC or better | D |  |  | 4 | 1 | 6
D7 |  | Temperature uniformity +/- 5oC (at 250oC) or better | D |  |  | 4 | 1 | 7
0 | Installation and Utilities |  |  |  |  | 4 | 1 | 7
M5 | Installation | Installation must be performed by qualifed engineers at CPI facility | M |  |  | 5 | 1 | 7
M6 | Site Requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment | M |  |  | 6 | 1 | 7
D8 |  | Ovens should be able to sit on standard laboratory bench, depth 600 mm | D |  |  | 6 | 1 | 8
M7 | Electrical Power | The system should operate from a standard UK mains supply (240Vac, 50Hz, single phase) | M |  |  | 7 | 1 | 8
M8 | Other Utilities | The system will not require a 3 phase supply, cooling water, or cryogenic gases | M |  |  | 8 | 1 | 8
D9 |  | Versions of the ovens should be available designed so that they can operate under inert atmospheres (e.g. Nitrogen gas) | D |  |  | 8 | 1 | 9
0 | Software and Connectivity |  |  |  |  | 8 | 1 | 9
D10 | Connectivity | Computer Interface (e.g. ethernet, USB etc) | D |  |  | 8 | 1 | 10
0 | Regulatory, Support and Warranties |  |  |  |  | 8 | 1 | 10
M9 | CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided | M |  |  | 9 | 1 | 10
M10 | Manuals | Instrument and controller manuals must be supplied | M |  |  | 10 | 1 | 10
O2 | Training | Where appropriate full training should be available along with access to technical support | O |  |  | 10 | 2 | 10
M11 | Warranties | Warranty period should cover at least 18 months from installation | M |  |  | 11 | 2 | 10
M12 | Extended Warranty | Extended warranty of up to 5 years | M |  |  | 12 | 2 | 10
M13 | Service Support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. | M |  |  | 13 | 2 | 10
M14 |  | Details of preventative maintenance programmes required (if any) above Warranty requirements | M |  |  | 14 | 2 | 10
[6 empty] | 14 | 2 | 10
[6 empty] | 14 | 2 | 10
[6 empty] | 14 | 2 | 10
[6 empty] | 14 | 2 | 10
[6 empty] | 14 | 2 | 10
[6 empty] | 14 | 2 | 10
[6 empty] | 14 | 2 | 10
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com
Note on F6: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: CHANGE THIS IN ACCORDANCE WITH THE TENDER DOC