**Specification**

Contract length: 5 yr maintenance service preventative cover. Includes two service visits per annum, service parts, functional check tech support, call-outs, labour and travel.

Equipment: Solar Gold - Urodynamic Investigation System s/n 22710914

Location: King’s College Hospital, Willowfield Surgical OPD

Response time: Telephone please confirm, on-site engineer – please confirm

Out of hours – please advice

Call-outs fees: please confirm

Loan unit: please confirm

Water filter replacement included: frequency of sending the replacement of water filters to be confirmed

Coverage working week: please confirm

Cancellation terms: please confirm

Payment terms: please confirm

6 monthly full replacements of wetted parts by our Technical Engineer on site, as per the manufacturers’ guidelines. The contract covers the cost of all the parts plus the call-out charge for the Technician. Parts required to be changed are as follows:

Connection tube combination

Tubes (flushing tube, tube from water container to water filter and water container tube)

Silicone gaskets of the water container lid and the refill cap

Water filters

Flow resistors for 0.6 ml/min (set of 8) / 0.15ml/min (set of 24)

TNF-R pressure transducers

PPM also includes:

* software build upgrade
* testing pressure range of compressor
* catheter protocol check
* pH calibration check of ambulatory products
* full clean of all components and equipment