## Instructions
Invitation to Tender (ITT) for the Provision of 2 x Trace Screening Instruments (TSI) for Explosives
Dstl Ref: IPMO 0825.22
INVITATION TO TENDER (ITT) RESPONSE WORKBOOK - TECHNICAL (QUALITY)
GUIDANCE & NOTES ON COMPLETION OF THIS WORKBOOK
Name of Organisation Completing this workbook:
Contact Name:
Contact Email Address and Telephone Number:
GUIDANCE FOR COMPLETION Tenderers must complete this workbook in its entirety. Your ITT response will be deemed non-compliant and excluded from this procurement if you fail to complete this workbook in full. Additional rows may be added to the workbook where required and only where it is indicated that this is permissible.
Note on Project Funding: Bidders are advised that funding for this requirement is subject to confirmation from internal Dstl sources which will be requested once the evaluation of all proposals has been completed and a suitable value for money solution identified. Award of any contract is subject to funding confirmation.

## Technical Response
Dstl Further Competition for the Provision of two Trace Screening Instruments (TSI) for Explosives
Reference: IPMO 0825.22
INSTRUCTIONS FOR COMPLETION AND SCORING
Please provide a response in all yellow cells.
STATEMENT OF REQUIREMENT
1. Mandatory Criteria
All listed Mandatory Criteria (including the stated sub-criteria) are essential to this procurement. Bidders must confirm that they are able to meet each of the stated criteria and sub-criteria, providing sufficient explanation to demonstrate this compliance to the Authority. Where appropriate, evidence should be provided in support of the response.
A response which fails to fully confirm compliance with the relevant criteria / sub-critieria will be marked as 'Fail'. A mark of 'Fail' in respect of one or more response(s) will result in your proposal being excluded from this procurement process.
Criteria | Sub-Criteria | Scoring Methodology | Tenderer's Response
1. Technical Requirement and Deliverables (Mandatory)
Two Trace Screening Instruments (TSI) meeting the requirements set out below.
System | The Trace Screening instrument (TSI) must be a state-of-the-art, computer controlled, rapid screening system, with the ability to analyse swabs taken from a wide range of materials for the presence of trace explosive compounds. The system must be mains powered and capable of being left on and running at all times. | PASS/FAIL – supplier to demonstrate that instrument meets these requirements.
The TSI will consist of:
· a tool that is capable of holding a swab firmly, whilst swabbing packaging, clothing, Tyvek suits, as well as any metal or plastic consumables;
· a system allowing for the analysis of positive and negative ions simultaneously;
· an alarm to alert the user to the presence of explosives; and,
· a sampling inlet to the instrument which is accessible for the user to clean and maintain.
System is designed such that:
· after a positive result, the instrument must be able to demonstrate it is clear of explosive contamination, i.e. no carryover;
· components are easily accessed for cleaning to remove contamination;
· sampling tool can be cleaned daily, with organic solvent to prevent explosive contamination build up;
· swabs are reusable, such that they can be analysed immediately prior to sampling to allow demonstration of being clean before;
· be programmable to run maintenance overnight, or during laboratory downtime; and,
· analysis is rapid for screening, ideally 15 seconds after swab insertion.
The TSI shall be shown to meet all essential specifications within this document.
Method | Instrument must be able to detect 5 ng/uL for nitroglycerine (NG), pentaerythritol tetranitrate (PETN), cyclotrimethylene trinitramine (RDX) and trinitrotoluene (TNT) in ethyl acetate, or better when directly applied to a swab. | PASS/FAIL – supplier to demonstrate that instrument meets these requirements.
The TSI must be able to store multiple methods for analysis of different explosive materials or cleaning procedures.
The list of target explosives, i.e. those which cause the system to alarm, must be able to be altered by the laboratory.
PC | The TSI must run Windows 10, or an equivalent operating system of adequate CPU resources, and utilise a Graphical User Interface (GUI) that gives an overview of analytes found, as well allowing the user to review all aspects of the results as required. | PASS/FAIL – supplier to demonstrate that instrument meets this requirement.
The TSI must employ an integrated touchscreen with the ability to easily interrogate the results on the system, i.e. not download onto a computer. Interrogation includes review of the data which has caused the system to alarm.
The TSI must have sufficient memory to store the results of 10,000 samples. Although sample results will be reviewed on screen in real time, they must be retrievable in full for scrutiny if required.
Software Functionality | The TSI operating software must provide a function to save and lock analytical parameters in a method. | PASS/FAIL
The software must be able to be updated without connecting the instrument to the internet.
Internet Connectivity | Item should not be fitted with WiFi or Bluetooth, or must have an option to permanently disabled (to allow use within a licensed explosives laboratory). | PASS/FAIL
Data Back-up | Ability to back up results and methods to external media for record retention. | PASS/FAIL
Equipment Integrity | Instrument must comply with relevant British safety standards, for example, Electrical safety. | PASS/FAIL – Supplier must state the appropriate International or Domestic Safety standard offered with their product.
Product size / weight | No larger than 1.0 m wide by 1.2 m deep | PASS/FAIL
Bench rated for 350 kg per metre
Status | The Instrument must be supplied new | PASS/FAIL
Warranty | The equipment must be provided with a minimum of 12 months warranty, to include parts, labour and annual service. | PASS/FAIL
2. Delivery / Installation Requirements
Delivery timescale | Instruments to be delivered by EOFY [End of March 2022] | PASS/FAIL
Delivery | Delivery is required to Dstl Porton Down, Salisbury, Wiltshire, SP4 0JQ. | PASS/FAIL
Installation | Instrument to be installed and commissioned by the manufacturer. | PASS/FAIL
3. Deliverable Acceptance Criteria
Installation | Instrument performance verified in-situ by completion of an accredited installation report. | PASS/FAIL
4. Post Delivery Service Requirements (Mandatory)
Product Support | Product 10-Year Support Package Required | PASS/FAIL
Manufacture Service/Preventative Maintenance can be carried out on-site.
Consumables | Consumables readily available in the UK from reliable manufacturers | PASS/FAIL
End-User Training | Training is available for at least 3 members of FEL staff, face to face and on site, within one month of delivery, complying with Covid restrictions as required. | PASS/FAIL
2. Desirable Criteria
The following Desirable Criteria and Additional Evaluation Critieria (including sub-criteria) form the technical scored element of this procurement. Bidders should provide a response against each critieria/sub-criteria, providing a full explanation and (where applicable) evidence in support of the response.
The Authority will score each response in accordance with the stated Scoring Methodology.
Please note: In the event that any response is scored as 'Fail' (as described in the Scoring Methodology), this will result in your proposal being excluded from this procurement.
Criteria | Sub-Criteria | Scoring Methodology | Tenderer's Response
5. Technical Requirement and Deliverables (Desirable)
Software Data Integrity | The software will meet current data integrity requirements, ALCOA, as a minimum, ensuring individual log ins, with audit trails, data retention and back up, time and date stamp work, ensure result files are linked to methods used to collect the data and cannot be edited at any access level without full audit trail. | 10 points if supplier can demonstrate full compliance
5 points if some aspects are met, including audit trails and individual log ins as a minimum.
0 points if no compliance can be demonstrated.
Maintenance | System should be easy to maintain by users, through including computer videos or graphical instructions for maintenance procedures. | 10 points for user access and Videos for all maintenance procedures
| The instrument must be fully accessible for users to clean and replace consumables. | 5 points for reduced access and instructions.
2 points for other maintenance help
System | A wand or similar means of inserting the swab for analysis without the need for the operator to touch the swab. | 5 points for the sampling tool being able to insert the swab for analysis
| The instrument has passed a recognised evaluation i.e. European Civil Aviation Conference (ECAC) Evaluation or equivalent. | 5 points for ECAC Evaluation or equivalent
| Wired connection to a printer or internal printer | 0 points (desirable but not enough to assign points)
Software Reporting | Software includes a flexible report template that is able to display all relevant information and can be tailored to user requirements. | 10 for fully flexible reporting options
5 points for an option to display a standard report
Analytical range | The ability to screen for a wider range of explosive compounds with a single scan is desirable. | 15 points for all explosives detectable
| (The instrument must be able to screen for 5 ng/uL NG, PETN, RDX and TNT as a minimum when applied directly to a swab.) | 10 points for minimum plus HMX, HMTD and TATP
| Other desirable explosives for detection are: | 5 points for minimum plus at least of one HMX, HMTD and TATP
| Organics Explosives: | 0 points for minimum only.
· Diazodinitrophenol (DDNP)
· Diethylene glycol dinitrate (DEGDN)
· Dinitrotoluene (DNT)
· Ethylene glycol dinitrate (EGDN)
· Erythritol tetranitrate (ETN)
· Cyclotetramethylene tetranitramine (HMX)
· Nitrobenzene (NB)
· Nitrotoluene (NT)
· Picric Acid
· Cylcotrimethylene trinitrosamine (R-salt)
· 2,4,6-Trinitrophenylmethyl nitramine (Tetryl)
Peroxide Explosives:
· 2,3-dimethyl-2,3-dinitrobutane (DMDNB)
· Hexamethylene triperoxide diamine (HMTD)
· Triacetone triperoxide (TATP)
6. Post Delivery Service Requirements (Desirable)
Product Warranty | Minimum 12 months warranty, to include parts, labour and annual service. | 5 points for 3 years warranty to include parts and labour
3 points for other warranty options
0 points if no warranty offered
Breakdown | Repair/breakdown call-out within 3 working days | 5 points– 3 working days
2 points – other
Support | Telephone support available to help resolve issues without call outs where possible. | 5 points – Supplier is able to demonstrate that this service is freely available to Dstl staff