Schedule A

### THE HEALTH AND SAFETY EXECUTIVE

Tender T26 15/16

STATEMENT OF REQUIREMENTS FOR THE PROVISION OF SUPPLY OF VACUUM STEAM AUTOCLAVE SYSTEM

1	THE HEALTH AND SAFETY EXECUTIVE (HSE)

1.1	The Health and Safety Executive (HSE) is a Crown non-departmental public body with specific statutory functions in relation to health and safety.  It is appointed by the Secretary of State for Work and Pensions and employs around 3500 staff including policy advisers, inspectors, technologists and scientific and medical advisers. HSE’s job is to prevent people being killed, injured or made ill by work.

  - HSE consists of a governing Board comprising of a Chair and nine non-executive members, all of whom are appointed by the Secretary of State for Work and Pensions after consultation with organisations representing employers, employees, local authorities and others, as appropriate.

1.3	HSE comprises various directorates and groups and is led by a senior management team.  The Health and Safety Laboratory (HSL) is part of HSE.  HSE works from over 30 locations throughout Great Britain.

2.	Health & Safety Laboratory

The Health and Safety Laboratory (HSL) is Britain's leading industrial health and safety facility. HSL supports HSE’s mission to protect people's health and safety by ensuring risks in the changing workplace are properly controlled.

HSL are inviting tenders for the supply of an autoclave steam sterilization system which will be used to process and ensure the safe disposal of hazardous biological waste.

The new unit will replace our current autoclave which is over 10 years old and recently has been experiencing more frequent breakdowns.

Most importantly the purchase is vital to ensure that biologically hazardous waste, generated both on site and from incidents, is steam treated in a safe and acceptable way.  Steam sterilization remains the ‘gold standard’ for such treatment.

3	BACKGROUND TO THE PROJECT

3.1	HSL’s Microbiology Team provides a range of analytical support and services to HSE and many external customers, which include the analysis and examination of potentially hazardous materials.  Even the most routine of microbiological testing may produce material that is contaminated with biological agents and requires disposal by effective means.  The Advisory Committee on Dangerous Pathogens (ACDP) stipulates that anything of a hazardous nature must be disposed of safely; ideally by steam sterilization facilities that are immediately available at the point the waste is generated.  The Microbiology Team has a large Priorclave steam autoclave for disposing of such waste, but the system is over 10 years old and has become unreliable in recent months.

In addition to the safe disposal of our own laboratory waste, the Microbiology team is regularly called upon to decontaminate many types of incident sample, for example, items that have been contaminated with biological residues.  Although this decontamination work may be done away from our main lab, the waste is then sealed and safely returned to the lab, where it is usually steam sterilized prior to end disposal.  In some cases, whole incident items may be treated in this way to render them safe for end disposal (providing they can be safely steam treated).

As well as the operational and health and safety priorities that require effective steam sterilization there are also business drivers for having a reliable, large volume steam sterilizer for the disposal of laboratory waste.

4		SPECIFICATION REQUIRED

The requirements are for an automated steam autoclave system that has sufficient capacity to accommodate laboratory waste (in purpose designed biohazard tins) to avoid unnecessary build-up of waste materials.  Specifically, the requirements are:

- A pressurised steam autoclave system with vacuum step, capable of sterilising hazardous waste and ensuring effective steam ingress in to packaged or bagged items

- Capable of running at 131°C for an hour under pressure – this cycle facility is a requirement for some of the hazardous materials that we must dispose of.

- Capable of running at least 1 cycle in a working day, but preferably 2.

- Capable of running for 450-500 hours per annum (average running time of around 2 hrs per day for between 45 and 50 weeks per year).

- Safety features such as interlocks for pressure and temperature to avoid operator opening system under dangerous pressures or temperatures.

- Must have a minimum square 350 litre capacity load space, capable of holding 2 bins of dimensions: 300mm x 300mm x 400mm in size, preferably with extra room for additional items. (current load space- 580mm x 600mm x 910mm)

- Current autoclave external dimensions are as follows 800mm x 1600mm x 1500mm (Approx.), we would hope to obtain a system around this size or smaller as there is little additional space within the installation area.

- Must have a printable recording system that can provide a print out of temperatures, times and pressures reached during the cycle; for safety audit purposes.

- Due to the location of the autoclave (small lab room) a site visit would be needed by any potential supplier to identify any problems that may arise during installation.  To formally arrange this please contact Mr Adam Clarke on: 01298 218492.

Installation, Commissioning and Training.

Tender to include itemised delivery, installation, and commissioning costs for the HSL Buxton site and provision of on-site training of 3-4 staff in the use of the new system.

Warranty.

Customary warranty or extended warranty to be offered. HSL would like to a achieve 24 month warranty period as standard for the instrument.

Other Services.

We would also like options for preventative maintenance and calibration services for the autoclave, to ensure its accurate thermal cycling performance going forward and to satisfy safety regulatory requirements for waste processing from Containment Level 2 and 3 laboratories.

Existing Equipment

We also wish tenderers to make an offer for the existing equipment, which is a Priorclave tactrol 2 (Type PS/RSC/EH350) , either for exchange or salvage, as part of this tender invitation.