Technical Specification – HSE/T3877

Hand-held portable equipment for measurement of compositions of metals

(both heavy and light elements) using XRF and/or LIBS analysis

| Area | Requirements | Explanatory notes |
| --- | --- | --- |
| Analysis capability | - Unit/s must be able to analyse compositions of metals and alloys such as low alloy steels, stainless steels, nickel alloys, copper alloys, titanium alloys. | Elemental analysis range to be specified (e.g. C, Al, S, V, Cr, Fe, Ni etc) |
|  | - Unit/s must be hand-held and portable, capable of use in both internal and external environments |  |
|  | - Unit, or units, must have ability to analyse elements in the range carbon to lead - Analysis of this elemental range may be achieved by one combined instrument or two individual instruments e.g LIBS for light elements and XRF for heavier elements. | Please provide quotes as appropriate for: - A single unit that covers the whole elemental range - Two separate units if required to meet the full elemental range - If only heavy elements can be analysed (XRF) please provide quote for that unit alone. |
| Accuracy/calibration | - Analysis capability to detect transition metals and elements from atomic number 13 and above with accuracy to 0.1 % or better. | Analysis accuracy to be provided with quotation. (if this is not achievable please state accuracy limits). |
|  | - Unit to be calibrated and supplied with suitable reference standards. Certificate of calibration required. |  |
| Portability | - Hand held capability required to allow analysis of large immovable samples. | Lightweight and portable such that it could be used on a scaffold/ladder |
| Data storage | - Memory capacity for multiple analyses (please state approx. number of analyses) |  |
|  | - PC compatible to allow downloading of results via Excel file or similar software. |  |
| Safety | - Safety features to protect the operator. | Please provide details of inherent safety features |
|  | - Protective and lockable carry case. |  |
| Support | - Annual safety and calibration check | Certificate of service and calibration required. |
|  | - Service and repair contract | Please provide details and cost |
| Additional features | - Please provide details of optional extras such as test stands. | Please cost separately |
|  | - Battery and charger must be included within the price |  |
| Additional features | Essential: - Water proof/water resistant (state IP rating). |  |
|  | Desirable; - Ability to analyse powders - Tilting display - Hot-swap battery system to prevent loss of data if batteries need replacing. |  |
|  | - Please advise of any additional hardware or consumable requirements such as argon/helium cylinders. |  |
| Warranty | - Please state warranty period |  |
| Lead time | - Please state best lead time from receipt of order |  |
| Delivery location |  | Delivery location is Buxton, Derbyshire, UK |

Bids will be evaluated and awarded on a Most Economically Advantageous Tender (MEAT) basis, including the ability to deliver before 31 March 2021.