## Drop Codes (hidden)
RoomClassification | IA | QualificationTest
Not Classified | Direct Impact | Yes
Grade A | Indirect Impact | No
Grade B | No Impact | n/a
Grade C
Grade D

## URS Document
User Requirement Specification |  | CPI Internal Tender Reference & Title:- | 1123 | Procurement of Multichannel Potentiostat
|  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender. | 40%
USER REQUIREMENT |  |  |  | Tenderer Response
Requirement codes |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as Desirable) |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter | Requirement |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
Channels | Must have at least 6 channels and be expandable further |  | M
Channels | Provide costs and options for additional channels |  | O
Channels | All channels can be run independantly |  | M
Range | Current range 1 Amp down to 10 nano Amps or less. |  | M
Experiments | Must be capable EIS, Cyclic Voltametry, DC Corrosion, Electrochemical Noise and Battery testing and have associated software modules. |  | M
Data fitting | EIS Fitting Software Provided |  | D
Acquistion rates | Data Aquasition at 1000 samples per second |  | D
Accessories and additional modulues | Provide options and costs of accessories including: Electrochemical cells (e.g. flat cells, corrosion cells and micro analytical cells). Additional channels Additional high current/voltage channels Biopotentiostats (either stand alone or integrated) Additional software licenses Low current modules/interfaces for pA range measurments |  | O
Consumables | Provide prices for Ag/AgCl electrodes |  | O
Freq range | Frequency Range 1MHz to 0.01 Hz. Please provide accuracies at different freuqencies and currents |  | M
Installation & Utilities
Installation & Utilities | Installation must be performed by qualifed engineers at CPI NetPark facility |  | M
Site requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment |  | O
Electrical | The system should operate from a standard UK mains supply (240Vac, 50Hz, single phase) |  | M
Other | The system will not require a 3 phase supply, cooling water, gases or cryogenic gases |  | M
Software and Connectivity
Connectivity | Computer Interface (e.g. ethernet, USB etc) |  | M
Regulatory, Support & Warranties
Safety | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided |  | M
Training | Full training should be provided along with access to technical support |  | D
Warranty | Warranty period should cover at least 18 months from installation |  | M
Extended warranty | Extended warranty of up to 5 years |  | D
Support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. |  | D
Servicess | Details of preventative maintenance programmes required (if any) above Warranty requirements |  | D
OFFICE USE ONLY - This section below this line is for CPI Evaluation and Internal USE only
|  | Name of Bidder to be inserted here | 1123 | Procurement of Multichannel Potentiostat
|  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com