Schedule A

### THE HEALTH AND SAFETY EXECUTIVE

Tender Ref T33 1516

STATEMENT OF REQUIREMENTS FOR THE PROVISION OF SUPPLY OF A

Permeation Diffusion Gas STANDARD Generator

1	THE HEALTH AND SAFETY EXECUTIVE (HSE)

1.1	The Health and Safety Executive (HSE) is a Crown non-departmental public body with specific statutory functions in relation to health and safety.  It is appointed by the Secretary of State for Work and Pensions and employs around 3500 staff including policy advisers, inspectors, technologists and scientific and medical advisers. HSE’s job is to prevent people being killed, injured or made ill by work.

  - HSE consists of a governing Board comprising of a Chair and nine non-executive members, all of whom are appointed by the Secretary of State for Work and Pensions after consultation with organisations representing employers, employees, local authorities and others, as appropriate.

1.3	HSE comprises various directorates and groups and is led by a senior management team.  The Health and Safety Laboratory (HSL) is part of HSE.  HSE works from over 30 locations throughout Great Britain.

2.	Health & Safety Laboratory

The Health and Safety Laboratory (HSL) is Britain's leading industrial health and safety facility. HSL supports HSE’s mission to protect people's health and safety by ensuring risks in the changing workplace are properly controlled.

3	BACKGROUND TO THE PROJECT

3.1	For the calibration/validation of air sampling methods, sample devices, sensors and the production of laboratory proficiency test samples, HSL requires the ability to produce stable standard gas atmospheres for a range of inorganic and organic gases.

4		OBJECTIVES

4.1.	HSL produces such stable standard gas atmospheres based upon procedures codified in the ISO 6145 series of standards.

- SPECIFICATION REQUIRED

The requirements are for:-

  - System based on diffusion/permeation process for the generation of gas standards as

codified in ISO 6145 part 8 and 10

  - Thermostatic oven covering the range of at least 30 to 100 ° C

Vendor to supply temperature set point resolution, stability and ranges and details of

temperature calibration protocol)

  - Vertical oven orientation allowing the use of diffusion cells

(Vendor to supply internal oven chamber dimensions, flow path construction materials and details/sizes of inlet and outlet gas connectors)

  - 1 x Mass Flow Controller carrier delivery or other metered flow system

(Vendor to supply details of mass flow ranges available and accuracy)

  - Desirable – Humidity control module to produce a relative humidity range of 20 – 90 %

(Vendor to supply details of achievable range, stability and set point resolution)

  - Desirable – RS-485 communications module for optional remote PC control via HSL software written in Labview, source code to be supplied.

  - Desirable – Modular system – e.g. extra chambers for multicomponent gas generation, dilution modules for secondary dilution

Installation, Commissioning and Training.  Tender to include itemised delivery, installation, and commissioning costs for the HSL Buxton site and provision of on-site training of 3-4 staff in the use of the new system.

Warranty.  Customary warranty or extended warranty to be offered. HSL would like to a achieve 24 month warranty period as standard for the instrument.

Other Services.

We would also like to look at preventative maintenance and calibration services