## Drop Codes (hidden)
RoomClassification | IA | QualificationTest
Not Classified | Direct Impact | Yes
Grade A | Indirect Impact | No
Grade B | No Impact | n/a
Grade C
Grade D

## Humidity Chambers
User Requirement Specification |  |  | CPI Internal Tender Reference & Title:- | 1132 | Framework for the Supply, Installation and Commissioning of Autogravimetric Analysers (Auto TGA)
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 40%
USER REQUIREMENT |  |  |  |  | Tenderer Response
Requirement codes |  |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory |  |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  | Requirement |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
| Instrument principles | A thermogravimetric analyser (TGA) capable of precisely measuring mass change in materials as a function of time and temperature with automated sample loading, testing and unloading |  | M
|  | Sample mass changes measured using a high sensitivity microbalance |  | M
|  | Compatible with measuring wide range of organic materials (e.g. polymers), inorganic oxides, metals and composites. Please state potential incompatible materials and at what temperature. |  | M
|  | Able to measure solid, powder and liquid samples |  | M
|  | Programmable, forced cooling achieved through integrated chiller unit |  | D
Technical Specification
| Instrument size | Benchtop unit |  | D
| Temperature range | 25 to 1200 °C |  | D
| Temperature accuracy & repeatability | Accuracy ± 1°C, repeatability ± 0.1°C |  | D
| Heating rates | Programmable heating 0.1 to 500°C/minute |  | D
| Mass capacity | 0.5 g |  | D
| Weighing resolution | 0.1 ug |  | D
| Weighing precision | ±0.01 % |  | D
| Volatile samples | Able to measure volatile samples in the automated system without loss of volatiles whilst awaiting test (sample may be sitting in the instrument 8 hours before test) |  | M
| Programmable autosampler | Able to programme to run upto 20 sequential samples in automated mode |  | M
| Static charge | Static charge dissipation included |  | D
| Sample pans | Sets of platinum and ceramic sample pans included. Disposable liners also included. Costs and part numbers for replacements listed. |  | M
| Atmosphere | Ability to run in air or under an Inert (nitrogen) atmosphere |  | M
| Standards | Set of appropriate standards and references included |  | M
| Spares kit | A set of typical spares included (please specify contents) |  | D
| Consummables | Please provide a list of standard consummables and their costs that maybe required when using the instrument. |  | O
Installation & Utilities
| Delivery | Please indicate delivery timescale from receipt of order. |  | M
| Installation | Installation must be performed by qualifed engineers at CPI NetPark facility |  | S
| Site requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment |  | M
| Electrical Power | The system should operate from single phase (13A or below) supply |  | D
Software and Connectivity
| Connectivity | Computer Interface (e.g. ethernet, USB etc) |  | D
| Instrument Software & Hardware | System specific software and fittings should be supplied with the instrument |  | M
| Computer hardware | Suitable PC imcluded |  | M
| Integration | Capabilities for integration to laboratory informatics systems should be detailed |  | D
| Software | Imaging in colour |  | D
| Software | The software should support instrument connectivity via TCP/IP, setup and operation, data collection, data processing, and export to file formats such as .csv , .xls and PDF. |  | D
| Software | Multiuser software license is desired to enable remote post test data analysis |  | D
| Future proofing | The system should have the ability to be upgraded with external accessories if / as required. Please provide details of potential upgrades. |  | D
Regulatory, Support & Warranties
| CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided |  | S
| Training | Full hardware and software training programmes should be provided along with access to application specific technical support |  | D
| Warranties | Warranty period should cover 36 months from installation |  | D
| Service support | On site support within 5 hours of requesting assistance. Please provide details of expected call-out charges. |  | D
| Service support | Details of preventative maintenance programmes required (if any) above Warranty requirements |  | D
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com