T35 1415 Ultra High Performance Liquid Chromatography (UPLC) with Photodiode Array (PDA) and Electro Chemical Detectors (EC)

- Introduction

The Health and Safety Laboratory (HSL) is Britain's leading industrial health and safety facility. Operating as an agency of the Health and Safety Executive (HSE), HSL supports HSE’s mission to protect people's health and safety by ensuring risks in the changing workplace are properly controlled.

HSL is inviting tenders for the supply of an Ultra High Performance Liquid Chromatograph (UPLC) with Photodiode Array (PDA) and Electro Chemical (EC) detectors. The new instrument will replace two existing HPLC instruments both of which are over 20 years old.

Current instrument configuration is Waters 600 pump unit coupled with Waters 717 water sampler and Waters 996 Photo Diode Array using a Coulochem II coupled with Thermofisher / Dionex 5010A Cell and ESA Guardcell this configuration produces the following analysis.

UPLC - PDA - EC is used at HSL to carry out quantitative analysis of organic compounds in solution, mainly monomeric and oligomeric isocyanates. Ideally, the requirement is for an analytical reproducibility of 1 % (or better).

The instrument will also be used to carry out quality assurance testing of proficiency testing (PT) samples prepared in the laboratory.  The PT samples require replicate batches of up to 75 samples to be analysed with a co-efficient of variation of less than 1.5%.

The new instrument should be bench top and have a footprint less than 850mm x 850mm (not including PC and VDU).  The instrument will handle about 3000 samples per annum, which equates to a running time of about 500 hours.

The instrument should be controlled, and data captured, by PC-based software using the Windows operating system. Please provide a separate price at schedule D for the PC and peripherals.

HSL are interested in receiving pricing for a new UPLC system, and may also consider a refurbished system not less than 3 years old. Ideally, HSL would like to trade-in the two existing HPLC instruments towards the new UPLC instrument. Exchange or removal of the older instruments may or will need to be delayed until after the new instrument is established and performing.

For the purposes of this invitation to tender it is necessary for HSL to be open and fair to possible sources that believe they can meet these requirements and may submit a tender.

Note to Tenderers:

The Selected Tender / supplier will be asked to run some of HSL’s samples on the equipment they propose to supply, to confirm sensitivity, stability and performance. HSL will not award a contract for supply until HSL are satisfied the instrument will perform as required.

- Specification

Instrument UHPLC / PDA / EC

Pump

- Capable of running HPLC and UPLC methods. The operating pressure must reach up to at least 1000 bar and have user selectable upper and lower pressure limits.

- Capable of operating with reverse phase and normal phase eluents.

- Should be a quaternary pump capable of blending up to four solvents simultaneously in any proportion.

- Should have an integrated degasser

- Should support flow ramping.

- Should be able to demonstrate that base line noise/ripple has been minimised.

- Support an automatic prime feature, to automatically prime the solvent lines

- Be able to deliver gradients with different curve shapes.

- Intake valve, plungers, gradient proportioning valve and seals should be accessible to users to perform maintenance without assistance from a service engineer.

Autosampler

- Accommodate at least 100 standard 2 ml vials

- Should be UPLC compatible, i.e. handle pressures up to at least 1000 bar

- Able to control temperature of sample compartment between 4 °C and 40 °C

- Precision equal or better than 0.5% RSD from 0.5 µl

- Injection volume range of at least 0.1 to 50 µl

- Capable of operating with reverse phase and normal phase eluents.

Column Management

- Be UPLC compatible ie for pressures of up to at least 1000 bar

- The column heater compartment will be able to support columns of 1.0 mm to 4.6 mm ID and up to at least 150 mm in length plus guard column

- The column manager compartment should cover the temperature range 4°C to 90°C

- Ideally two or more columns with program controlled column switching should be supported

Photo Diode Array Detector

- Designed for use with UPLC systems

- Cover a wavelength range of at least 190 to 800 nm

- Have simultaneous 3D and 2D operation

- Have data rates of at least 100 Hz

- Optical resolution should be in the order of 1 nm

- Should have a low cell volume preferably < 15 µl

Electrochemical detector

- Designed for use with UPLC systems

- Capable of working in line with a PDA detector

- Have a DC mode electrochemical cell designed for use with UPLC pressures and have a porous graphite electrode capable of operating at voltages of least 1000 mV

- Have an electrochemical guard cell placed between the auto sampler and column for use with UPLC pressures and capable of operating at voltages of at least 1500 mV or be able to demonstrate a clean mobile phase when analysing HSL test samples

System

- Where multiple suppliers are proposed, compatibility of systems must be demonstrated.  There should be two-way communication between the PC and all components with full PC control of those components where possible.  Any manual controls or non 2-way communication must be explicitly mentioned

- Instrument

The supply of the instrument will include all components hardware and software and appropriate electrical gas and data connections.

- Installation, Commissioning, Training

Tender to include installation commissioning and delivery to HSL Buxton Derbyshire and on-site training.

- Warranty and Documentation

Customary warranty and/or extended warranty to be offered.

All user manuals and copies of software and manuals should be received by HSL with the system hardware.

On warranty, HSL would like to achieve 24 month warranty period for the instrument as standard.

- Information Technology

HSL has a separate Facilities Management IT Provider whom are responsible for the whole network of the laboratory.  In that regard, they have a clear responsibility for ‘impact testing’ new IT equipment onto HSL’s IT network. The laboratory’s FM IT Provider sets a minimum requirement on all new equipment.

It is therefore a requirement that all Tenderers should specify in the fullest detail possible, the makers name and specification etc of all IT equipment (hardware, software etc.) that they propose to supply including details of the proposed operating system. Details of any associated IT peripherals and software should also be supplied. It would greatly assist if separate itemised costs for all of the associated IT components (e.g. PC, LCD, and Printer) could be included within the tender response, should these be available.

In addition, please provide warranty information, stating whether it is standard warranty (return to base) or onsite warranty (next business day), together with the length of warranty that will be supplied with the proposed equipment. Please also provide details of exactly how these warranties shall be activated and/or requested, should any problems occur.

PC Minimum Specification:

Processor: Intel i5 (or equivalent)

Memory: 4GB

Hard disk: 500GB

Display: DVI capable of 1280 x 800 resolution

Operating System: Windows 7 Professional

Connectivity: Network adapter (Ethernet RJ-45)

The PC is considered as part of the 'system' we buy and should be specifically included in any warranty agreement offered by the tenderer. We would expect that any problems or repairs with the software and/or hardware that forms part of the 'system' should be resolved by the supplier.

- Software

Prior to placing any order, HSL requires assurance from the vendor that the proposed software contained within the order does not require administration rights to run and operate correctly post installation. This is because HSL is unable to run and operate software in MS Windows that needs to have administration rights before it will run and operate correctly post installation. Therefore we wish the vendor to:

- Confirm that the software does not require administration rights to run; or

- If you are unable to provide a version of software that does not run and operate correctly without “administration rights”, please advise what you can do to resolve this.

Note that in the event of an order being placed, and it being found that “administration rights” are required to run and operate the software, or, alternatively, that you do not notify HSL's Purchasing Procurement Unit that the software needs administration rights in order to run and operate correctly, HSL reserves the right to cancel the order and to place an order with another supplier and to seek a refund of any payments that have been made against the purchase order.