# Draft goods information

# Uniaxial mechanical testing unit

- Parameters

- The Supplier shall provide a tension-compression materials test rig that provides the following parameters:

  - Frame: The load frame must consist of two columns so that the system can be adapted with an environmental chamber in the future.

  - Testing regimes: The system must allow static, monotonic testing and programmable cycling capability.

  - Max load (tension): The maximum load must be 5 kN or greater, but with strong control and accuracy at low forces. Load cells will have a 16-bit resolution. Load cells (and any switchable load ranges) will be calibrated to EN ISO 7500-1 standard.

  - Cyclic testing and frequency:  The system must be capable of tension-tension, compression-compression and tension-compression cyclic testing. At 1 Hz the system must be capable of 2 mm displacement.

  - Analogue outputs: There must be 2 or more 0-10 V analogue outputs for load and displacement.

  - Analogue inputs: There must be one or more analogue 0-10 V input for extensometry.

- Extensometry: No explicit extensometry is to be provided by the supplier, but direct view of the sample is essential as the system will be used alongside Digital Image Correlation (DIC) systems for measurements of strain.

- The system should be capable of excellent control:

  - Minimum displacement of 20 µm or better; associated accuracy of 20 µm or better.

  - Minimum force of 500 N or better.

- All parts must have CE certification markings on them to ensure conformity with the health, safety, and environmental protection standards for products sold within the European Economic Area.

- Rig height: < 3.3 m.

- Floor force: < 2 tonnes per m2.

- Sample dimensions

- Flat dogbone samples. Total length 50mm, cross head dimensions 10 x 12 mm, gauge length 3 x 26 mm, thickness 1 mm.

- Cylindrical samples with threaded ends. Total length 27 mm, threaded dimensions 3 mm diameter x 6 mm, gauge length 2 mm diameter x 15 mm.

- Required features

- Control system: Comprehensive control system capable of complicated and programmable testing patterns. Fully user programmable waveforms can be produced within the software (profile builder) or imported from Excel or other programmes.

- Tensile grips applicable to typical specimen types and sizes mentioned. The grips will be capable of adjustable gripping forces. Self-aligning grips preferable but not essential.

- The system will be capable of being aligned to within the 5% in accordance with ASTM E1012 as required and be delivered as such. An associated alignment calibration certificate will be supplied on delivery.

- Offer to include fully strain-gauged alignment specimen. Standard type specimen. The system must allow a full specimen alignment routine by the operator to enable true uniaxial stress states.

- Desirable features

- Grips qualified up to 800 °C.

- Inclusion of extensometers as an extra. Due to the system planned to be used in conjunction with DIC, we are most interested in ‘standard’ clip-on extensometers but will consider other options.

- Additional 0-10 V outputs.

- Additional analogue inputs for extensometry.

- Training requirements

- Basic training to include use of software for best design and interpretation of testing. Software manuals to be included.

- Comprehensive training in the use of the rig, control software, maintenance and safety features.

- Servicing

- The supplier must be available for on-site servicing and maintenance requirements as necessary.

- Supply requirements

- Delivery and installation of equipment by 31st March 2019.

- The Supplier shall deliver the goods to the Purchaser’s site under Incoterms ‘Delivered Duty Paid’.

- The Supplier shall ensure that the goods are appropriately packaged.

- The Supplier shall offload, install and fully commission the system at the site in the precise location of the Purchaser’s choosing