## Sheet1
| User Requirement Specification | CPI Tender Reference & Title:- | 1641 | Laboratory Tangential Flow Filtration Apparatus
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column C, of which the total Quality element score of this sheet will equate to value in cell F6 | Scoring Criteria 0 = no response/answer not relevant to requirement 1 = limited response to requirements 2 = less than adequate, filfils some of the basic requirements 3 = adequate response, fulfils all basic requirements 4 = good response, fulfils all requirements to a high standard 5 = excellent response, fulfils requirements to an exceptionally high standard |  | The total percentage available which will contribute towards the Quality scoring element of this tender | 40%
USER REQUIREMENT |  |  |  | TENDERER RESPONSE
Requirement Codes |  |  |  | Guidance
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as desirable) |  |  | Tenderers must provide responses within the following columns; 1. Column F - please provide details of your equipment functionality in relation to CPI requirements specified in column E 2. Column G - please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response |  | Y
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  |  | N
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
# | Parameter | Requirement | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N) | N
System Performance
M1 | General Requirements | A single system that has the ability to perform the following sequence of controlled processes; constant feed followed by diafiltration followed by concentration. | M |  |  | 1 | 0 | 0
M2 |  | A single fully integrated tangential flow filtration apparatus | M |  |  | 2 | 0 | 0
M3 |  | Interchangable with a range of filtration media e.g. PES, mPES, cellulose | M
M4 |  | Tubing connectors and spares kit, suitable tubing, spare tubing | M |  |  | 4 | 0 | 0
M5 | Operation | Compatible with a range of organic solvents including ethanol and methanol | M |  |  | 5 | 0 | 0
M6 |  | Have fully automated back pressure valve to control retentate pressure controlable using the control software via the associated computer | M |  |  | 6 | 0 | 0
M7 |  | Supplied with main fitrate pump and two auxillary pumps to control diafiltration | M |  |  | 7 | 0 | 0
M8 |  | Main filtrate pump must be peristaltic pump to prevent contact of metallic moving parts with liquid | M
M9 |  | Should be a hollow fibre filtration device operable with hollow fibre filters | M
M10 |  | Should be a flat sheet filtration device operable with flat sheet filters | M
M11 |  | Must be supplied with a range of flat sheet cartridges and hollow fibre filters to allow both commissioning and life in service, size to be specified at time of ordering | M
M12 | Automation and Control | Must be supplied with computer/laptop with control software | M
M13 |  | Must have balances for permeate and retentation control | M |  |  | 13 | 0 | 0
M14 |  | Must have auto-shutdown capability based on weight | M |  |  | 14 | 0 | 0
0 | Installation and Utilities |  |  |  |  | 14 | 0 | 0
M15 | Installation | Installation must be performed by qualifed engineers at CPI NetPark facility | M |  |  | 15 | 0 | 0
M16 | Site Requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment | M |  |  | 16 | 0 | 0
M17 | Electrical Power | The system should operate from a standard UK mains supply (240Vac, 50Hz, single phase) | M |  |  | 17 | 0 | 0
M18 | Other Utilities | The system will not require a 3 phase supply, cooling water, gases or cryogenic gases | M |  |  | 18 | 0 | 0
M19 |  | The system will be bench mounted | M |  |  | 19 | 0 | 0
0 | Software and Connectivity |  |  |  |  | 19 | 0 | 0
M20 | Connectivity | Computer Interface | M |  |  | 20 | 0 | 0
O1 |  | Touch screen glove compatible | O |  |  | 20 | 1 | 0
O |  | USB connectivity | O
O3 |  | WiFi connectivity | O |  |  | 20 | 3 | 0
0 | Regulatory, Support and Warranties |  |  |  |  | 20 | 3 | 0
M21 | CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided | M |  |  | 21 | 3 | 0
D1 | Training | Full training should be provided along with access to technical support | D |  |  | 21 | 3 | 1
M22 | Warranties | Warranty period should cover at least 18 months from installation | M |  |  | 22 | 3 | 1
D2 | Extended Warranty | Extended warranty of up to 5 years | D |  |  | 22 | 3 | 2
D3 | Service Support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. | D |  |  | 22 | 3 | 3
D4 |  | Details of preventative maintenance programmes required (if any) above Warranty requirements | D |  |  | 22 | 3 | 4
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com