## Drop Codes (hidden)
RoomClassification | IA | QualificationTest
Not Classified | Direct Impact | Yes
Grade A | Indirect Impact | No
Grade B | No Impact | n/a
Grade C
Grade D

## Lot 1 - HTPD
User Requirement Specification |  | CPI Internal Tender Reference & Title:- | 1088 | Lot 1 - High Throughput Scalable Fermentation System
The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 40%
USER REQUIREMENT |  |  |  | Tenderer Response
Requirement codes |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
| System | A bioreactor system with capabilities for microbial and mammalian cells capable of parallel culture of at least 12 reactors simultaneously | M
| Volumes | Should be capable of culturing volumes in the range 150-500ml | M
| Monitoring | Temperature, pH and DOT for the duration of processes. Software should allow the comparison of all cultures being run. | M
| Control | Independent temperature control, pH control, DOT control, sterile feed additions. | M
| Sterility | Must be fully contained and demonstrably free from contaminations | M
| Gases | Should be sparge gas option for ease of scale up.. Suitable for maintaining the correct DOT and pH of the system. Provision for air, CO2 and O2. | M
| Scalable | Generic data and case studies showing system is scalable | M
| Shaker speed | Impeller speed of 10-1000rpm to allow sufficient aeration of cultures | M
| Automation | Automation of routine culture maintainence in order to reduce resource footprint including automated cell counting whether part of the same system or through integration with other equipment | M
| Support | Must be service engineers available for help/advice in the event of the system malfunctioning. A 3 year maintenance agreement would desirable. | M
| Training | Full operator training on the safe operation of the equipment and associated software routine maintenance and troubleshooting | M
| Connectivity | Costs for OPC/LIMS connection should be included. | M
| Software | If standalone software is required to analyse run data, costs for at least 3 licenses should be included. | D
| Consumables | The consumable costs for 5 complete runs should be included. | M
Installation & Utilities
| Installation | Costs for installation at NBMC, Darlington should be included | M
| Site requirements | Size requirements should be detailed as should utilies required for function. | M
| Electrical Power | Requirements should be detailed | M
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com

## Lot 2 - Mammalian Suspension
User Requirement Specification |  | CPI Internal Tender Reference & Title:- | 1088 | Lot 2 - Mammalian Bioreactors (Laboratory And Pilot Scale)
The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 40%
USER REQUIREMENT |  |  |  | Tenderer Response
Requirement codes |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
| System | A disposable cell bag system aerated by rocking motion. The system should be mobile on skids or a trolley and movable to different processing areas | M
| Scale | The system should have interchangeable platforms suitable for bag sizes from 5 to 50L (working volumes from about 2-20L). | M
| Temperature control | Platforms should be heated with inbuilt temperature sensor | M
| Rocker speed (rpm) | 7 to 45 | M
| Rock angle (o) | 3 to11 | M
| Gas mixer/controller | Ability to mix air with CO2 (2-10%) and deliver to the bag at flow rates concomitant with successful culture of mammalian cells. | M
|  | Must allow for the monitoring and control of temperature, pH and DOT (via optical sensors in bag). Data should be logged throughout the course of the cell culture. | M
|  | Gas mixing could occur in the control module | D
| Support | Must be evidence of locally based service engineers for advice/support and site visits | M
| Training | Full operator training on the safe operation of the equipment, routine maintenance and troubleshooting | M
| Compatability | Must demonstrate equivalent performance and scalability with existing Sartorius Rocker. Rationale and/or data should be given to support this. | M
| IQ/OQ | Qualification of the system as fit for purpose with relevant data pack must be provided following installation | M
| Consumables | The cost of consumables for 5 controlled runs with a 10L working volume should be included | M
| Software | If standalone software is required to analyse run data, costs for at least 3 licenses should be included. | D
Installation & Utilities
| Installation | Costs for installation at NBMC, Darlington should be included | M
| Site requirements | Size requirements should be detailed as should utilies required for function. | M
| Electrical Power | Requirements should be detailed | M
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com

## Lot 2 - Mammalian Lab Reactors
User Requirement Specification |  | CPI Internal Tender Reference & Title:- | 1088 | Lot 2 - Mammalian Bioreactors (Laboratory And Pilot Scale)
The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 40%
USER REQUIREMENT |  |  |  | Tenderer Response
Requirement codes |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
| System | Laboratory-scale benchtop system with independent controllers. System should be expandable should more vessels be required in the future. | M
| Vessels | 3 x10L reusable glass vessels (with the option of using disposable vessels). Dimension of vessels should fit the facility autoclave (dimensions of which are W=650cm, L=1340cm, H=1000cm). Aspect ratio of vessels should be equivalent to that of 200L disposable bioreactor. | M
| Vessel type and operation | Stirred tank operated in batch or fed-batch mode. Impeller must be equivalent to the 200L disposable vessel. | M
| Gas flows | Air, N2, CO2, O2 via sparger line. Separate line for gas to vessel headspace. | M
| Sterilisation | By autoclave or gamma irradiated disposable vessels | M
| Standard probes/sensors | Temperature, pH and DO. Provision for amperometric or optical probes. Probes to be supplied with vessel. | M
| Ports | Ports in vessel headplate for temperature, pH and DO probes. At least 2 spare ports for novel probes/sensors (biomass, IR, RI, Raman). Additional ports for feeds, base and antifoam addition. Provision of septum for bolus addition of feeds. Dip tube for off-line temperature monitor for calibration. | M
| Sampling | Ability to take samples for off line, in-line or at-line analysis. | M
| Controller | PLC based controller that can be used in conjunction with SCADA. Must be up to 12 configurable inputs/outputs available. | M
|  | Critical process parameters (temperature, pH and DOT) must be measured, controlled and logged for the duration of the cell culture process. | M
|  | PID control loops for temperature, pH and DOT are expected. DOT controlled by gas cascade. | M
|  | Mass flow controllers for N2, air, O2 and CO2 must be included. Flow rates at 10L scale must be demonstrably suitable for a wide range of mammalian cells (e.g. CHO, NS0, MDCK). | M
|  | Must allow for the stable, precise mixes of 2 or more gases through the sparger. | M
|  | Gas flows (N2, air, O2 and CO2) must be measured, controlled and logged for the duration of the cell culture process. | M
|  | Provision for rotameters for lower gas flows (e.g. if CO2 bleed required). | M
|  | Stirrer speed must be logged and in the range 20-800 rpm. | M
| Novel probes/sensors | Able to integrate biomass probes with controller and/or SCADA | M
| Controller alarms | For temperature, pH & DOT. Remote alarm alerts are desirable. | M
| Advanced control | Time or event based parameter changes, for example set point changes, real time process control, complex feed strategies | M
| LMS connectivity | 2 way data communication between the controller and LMS | M
| Support | Must be evidence of locally based service engineers. A 3 year maintenance agreement should be included | M
| Training | Full operator training on the safe operation of the equipment, routine maintenance, troubleshooting and advanced course in associated software | M
| IQ/OQ | Qualification of the system as fit for purpose with relevant data pack must be provided following installation | M
| Compatability with existing systems | Must demonstrate equivalent performance and scalability with existing Sartorius B-DCU/ 10L univessel systems. Rationale and/or data should be given to support this. | M
| Software | If standalone software is required to analyse run data, costs for at least 3 licenses should be included. | D
| Consumables | The cost of consumables for 10 bioreactor runs using glass vessels and/or 10 x 10L disposable vessels | M
| Spares | Critical spares to be included | M
Installation & Utilities
| Installation | Costs for installation at NBMC, Darlington should be included | M
| Site requirements | Size requirements should be detailed as should utilies required for function. | M
| Electrical Power | Requirements should be detailed | M
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com

## Lot 2- 200L Cell Bioreactor
User Requirement Specification |  | CPI Internal Tender Reference & Title:- | 1088 | Lot 2 - Mammalian Bioreactors (Laboratory And Pilot Scale)
The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 40%
USER REQUIREMENT |  |  |  | Tenderer Response
Requirement codes |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
| System | A single-use bioreactor shell with an integrated controller. The system should be mobile and easily wheeled to different processing areas. | M
| Scale | Ability to run single-use disposable vessel up to 200L. The bag should have an operating volume between 50 and 200L. Aspect ratio of vessel should be equivalent to laboratory-scale systems. | M
| Vessel type and operation | Stirred tank operated in batch and fed-batch mode. Impeller must be equivalent to those used in the 10L vessel. | M
| Gas flows | Air, N2, CO2, O2 via sparger line. Separate gas line to headspace. | M
| Sterilisation | Single use, presterilised bag | M
| Standard probes | Temperature, pH and DO. Probes should be integrated with the disposable bag. | M
| Ports | Ports for temperature, pH and DO probes. Spare ports for novel probes (biomass, IR, RI, Raman). Additional ports for feeds, base and antifoam addition. Dip tube for off-line temperature monitor for calibration. Surface and sub-surface additions should be possible. | M
| Sampling | Ability to take samples for off line, in-line or at-line analysis | M
| Controller | PLC based controller. Must be up to 12 inputs/outputs available. | M
|  | Critical process parameters (temperature, pH and DOT) must be measured, controlled and logged for the duration of the cell culture process. | M
|  | PID control loops for temperature, pH and DOT are expected. DOT controlled by gas cascade. | M
|  | Mass flow controllers for N2, air, O2 and CO2 must be included. Gas flow rates at 200L scale should be suitable for a range of mammalian cells (e.g. CHO, NS0, hybridoma) | M
|  | Must allow for the stable, precise mixing of 2 or more gases into sparger. | M
|  | Gas flows (N2, air, O2 and CO2) must be measured, controlled and logged for the duration of the cell culture process. | M
|  | Stirrer speed must be logged and up to approximately 200 rpm. | M
|  | A sufficient number pumps for acid, base, feeds are required. Gravimetric control should be possible (on continuous feeds for example). | M
|  | The controller should have OPC connectivity. | M
| Alarms | For temperature, pH, DOT, stirrer. Remote monitoring and alerts | M
| Advanced control | Time or event based parameter changes, real time process control, complex feed strategies | M
| LMS connectivity | 2 way data communication including integration with control system | M
| Training | Full operator training on the safe operation of the equipment and advanced course in any associated software | M
| Support | Must be evidence of locally based service engineers for advice/support and site visits | M
| Software | If standalone software is required to analyse run data, costs for at least 3 licenses should be included. | D
| Compatability | Must demonstrate equivalent performance and scalability with existing Sartorius 200L SUB. Rationale and/or data should be given to support this. | M
| IQ/OQ | Qualification of the system as fit for purpose with relevant data pack must be provided following installation | M
| Consumables | The consumable costs for 10 x 200L run must be estimated | M
Installation & Utilities
| Installation | Costs for installation at NBMC, Darlington should be included | M
| Site requirements | Size requirements should be detailed as should utilies required for function. | M
| Electrical Power | Requirements should be detailed | M
OFFICE USE ONLY - This section below this line is for CPI Evaluation and Internal USE only
| Notes for Evaluation Team: | Name of Bidder to be inserted here | 1088 | Lot 2 - Mammalian Bioreactors (Laboratory And Pilot Scale)
Please assess all of the answers given, with particular attention to the Mandatory answers provided, taking into account, the Desirable and Optional answers too, however, you do not need to score each individual answer, but based on the scoring range 0-5, just give one singluar score (default set at "5" in the red and yellow box below, and as we have to provide feedback to bidders, you must also provide comments on which section(s) above where you felt that they did not meet the requirements.
| Evaluation Panel | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com

## Lot 3- Adherent Cell
| User Requirement Specification |  | CPI Internal Tender Reference & Title:- | 1088 | Lot 3 - Mammalian Adherent Bed Reactor
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 40%
| USER REQUIREMENT |  |  |  | Tenderer Response
| Requirement codes |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
| M | Mandatory |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
| O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
| D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
| Parameter |  |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
|  | System | A single use bioreactor capable of producing product from high-density adherent cells. | M
|  | Scale | Capable of producing grammage amounts of product per month. Available culture area of at least 1m2. | M
|  | Vessel type | Perfusion based system with a suitable matrix for mammalian cells to adhere | M
|  | Sampling | Ability to take samples for off-line, in-line or at-line analysis | M
|  | Sterility and containment | Must be a contained, single use system. Demonstrably suitable for the manufacture of virus and virus like particles. Documented evidence of regulatory approval for planned activities | M
|  | Controller | Must be monitoring and automated control of temperature pH, perfusion and DOT. | M
|  |  | A biomass sensor would be desirable. | D
|  |  | Provision for air, CO2 and O2. The gas module should allow a suitable strategy for maintenance of a high cell density. | M
|  |  | Alarms to alert operators should deviations from set points occur. | M
|  | Data recording | Data must be able to be recorded and trends observed | M
|  | Consumables | Cost of consumables for 3 runs must be included | M
|  | Software | If standalone software is required to analyse run data, costs for at least 3 licenses should be included. | D
|  | Training | Full operator training on the safe operation of the equipment particularly if used for virus production | M
|  | Support | Must be evidence of locally based service engineers for advice/support and site visits | M
|  | IQ/OQ | Qualification of the system as fit for purpose with relevant data pack must be provided following installation | M
Installation & Utilities
|  | Installation | Costs for installation at NBMC, Darlington should be included | M
|  | Site requirements | Size requirements should be detailed as should utilies required for function. | M
|  | Electrical Power | Requirements should be detailed | M
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com

## Lot 4- Microbial Lab
User Requirement Specification |  | CPI Internal Tender Reference & Title:- | 1088 | Lot 4 – Microbial Fermentors (Laboratory And Pilot Scale)
The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 40%
USER REQUIREMENT |  |  |  | Tenderer Response
Requirement codes |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
| System | Laboratory-scale benchtop system with independent controllers. System should be expandable should more vessels be required in the future. | M
| Vessels and Scale | 3 x10L glass vessels to be provided. The option to run disposable vessels (2 – 10L) from the same controller is desirable. | M
| Vessel type and operation | Stirred tank operated in batch, fed batch or continuous mode | M
| Gas flows | Air, N2 and O2 via sparger line. | M
| Sampling | Ability to take samples for off-line, in-line or at-line analysis | M
| Sterilisation | By autoclave (if not single use) | M
| Controller | PLC based controller that can be used in conjunction with SCADA. Must be up to 12 configurable inputs/outputs available. | M
|  | Critical process parameters (temperature, pH and DOT) must be measured, controlled and logged for the duration of the cell culture process. | M
|  | PID control loops for temperature, pH and DOT are expected. DOT controlled by gas cascade. | M
|  | Mass flow controllers for N2, air, O2 must be included. Flow rates at 10L scale must be demonstrably suitable for a wide range of microbial cell types (e.g. E.coli, Saccharomyces) | M
|  | Must allow for the stable, precise mixes of 2 or more gases through the sparger. | M
|  | Gas flows (N2, air, O2) must be measured, controlled and logged for the duration of the cell culture process. | M
|  | Stirrer speed must be logged and up to 1000 rpm. | M
|  | Sufficient number of pumps for base, acid, antifoam and feed additions. Gravimetric control should be possible (on continuous feeds for example). The quoted system should include one, or more desireably two, gravimetrically controlled pumps with integrated balances. | M
|  | It is desirable that the pumps are compatible with different size silicone tubing and have appropiate output ranges to ensure maximum process flexibility and control. Could the pump output ranges (ml/min) please be provided. | D
|  | The controller should have OPC connectivity | M
| Advanced control | Time or event based parameter changes, real time process control, complex feed strategies | M
| Probes | Temperature, pH and DO. To be supplied with vessel. | M
| Novel probes | Provision for novel probes (biomass, IR, RI, Raman), and able to integrate biomass probes with controller and/or SCADA | M
| Pressure control | Measurement and control of reactor pressure would be desirable | D
| Measure, log and control | Temperature, pH, DO, gas flows, stirrer speed | M
| Sterile additions | Feeds, acid/base for pH control, antifoam (provision for continuous or bolus additions) | M
| Off gas analysis | Yes (O2, CO2) | M
| LMS connectivity | 2 way data communication including integration with control system | M
| Training | Full operator training on the safe operation of the equipment and advanced course in any associated software | M
| Support | Must be evidence of locally based service engineers for advice/support and site visits. It is desireable that the offered service contract should cover all peripheral equipment items (eg. off-gas analyser, balances). | M
| IQ/OQ | Qualification of the system as fit for purpose with relevant data pack must be provided following installation | M
| Consumables | The cost of consumables for 10 bioreactor runs using glass vessels and/or 10 x 10L disposable vessels | M
| Software | If standalone software is required to analyse run data, costs for at least 3 licenses should be included. | D
Installation & Utilities
| Installation | Costs for installation at NBMC, Darlington should be included | M
| Site requirements | Size requirements should be detailed as should utilies required for function. | M
| Electrical Power | Requirements should be detailed | M
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com

## Lot 4- Microbial Pilot
User Requirement Specification |  | CPI Internal Tender Reference & Title:- | 1088 | Lot 4 – Microbial Fermentors (Laboratory And Pilot Scale)
The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 40%
USER REQUIREMENT |  |  |  | Tenderer Response
Requirement codes |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
| System | A steam-in-place (SIP) fermentor with an integrated controller. Dedicated for microbial fermentations. The system should be mobile and easily wheeled to different processing areas. | M
| Scale and vessel | 1 x 100L (working volume) fermentor. Must be equivalent to the laboratory-scale bioreactors in terms of aspect ratio, probes and impeller types | M
| Vessel type | Stirred tank operated in batch or fed-batch mode. Options to operate as a continuous fermentor would be desirable | M
| Gas flows | Air, N2 and O2 via sparger line or headspace | M
| Sampling | Ability to take samples for off line, in-line or at-line analysis. At least one sample port should be resterilisable using steam | M
| Sterilisation | Steam-in-place using plant steam. A semi-automated sterilisation sequence would be desirable | M
| Controller | PLC based controller. Must be up to 12 configurable inputs/outputs available. | M
|  | Critical process parameters (temperature, pH and DOT) must be measured, controlled and logged for the duration of the cell culture process. | M
|  | PID control loops for temperature, pH and DOT are expected. | M
|  | Temperature control by electrical heating, steam only required for sterilisation. | M
|  | DOT controlled by gas cascade. | M
|  | Mass flow controllers for N2, air, O2 must be included. Flow rates at 100L scale must be demonstrably suitable for a wide range of microbial cells (e.g. E.coli, Saccharomyces). | M
|  | Must allow for the stable, precise mixes of 2 or more gases through the sparger. | M
|  | Gas flows (N2, air, O2) must be measured, controlled and logged for the duration of the cell culture process. | M
|  | Stirrer speed must be logged and in the range suitable for homogeneous microbial cultures. | M
|  | Sufficient number of pumps for base, acid, antifoam and feed additions. Gravimetric control should be possible (on continuous feeds for example). The quoted system should include one gravimetrically controlled pump with integrated scale. | M
|  | The controller should have OPC connectivity. | M
| Advanced control | Time or event based parameter changes, real time process control, complex feed strategies | M
| Probes/sensors | Temperature, pH and DO. To be supplied with vessel. | M
| Novel probes/sensors | Provision for novel probes (biomass, IR, RI, Raman), and able to integrate biomass probes with controller and/or SCADA | M
| Measure, log and control | Temperature, pH, DO, gas flows, stirrer speed | M
| Sterile additions | Feeds, base for pH control, antifoam (provision for continuous or bolus additions) | M
| Off gas analysis | Yes (O2 and CO2) | M
| LMS connectivity | 2-way data communication including integration with control system | M
| Training | Full operator training on the safe operation of the equipment and advanced course in any associated software | M
| Support | Must be evidence of locally based service engineers for advice/support and site visits. It is desireable that the offered service contract should cover all peripheral equipment items (eg. off-gas analyser, balances). | M
| IQ/OQ | Qualification of the system as fit for purpose with relevant data pack must be provided following installation | M
| Consumables | The consumable cost for a 100L run must be estimated | M
| Software | If standalone software is required to analyse run data, costs for at least 3 licenses should be included. | D
Installation & Utilities
| Installation | Costs for installation at NBMC, Darlington should be included | M
| Site requirements | Size requirements should be detailed as should utilies required for function. | M
| Electrical Power | Requirements should be detailed | M
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com

## Lot 5 - Analysers
| User Requirement Specification |  | CPI Internal Tender Reference & Title:- | 1088 | Lot 5 – Fermentation and Cell Culture Analysis
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 40%
| USER REQUIREMENT |  |  |  | Tenderer Response
| Requirement codes |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
| M | Mandatory |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
| O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
| D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
| Parameter |  |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
|  | System | Benchtop metabolite analysers for in process samples from bioreactors and fermentors from microscale to large scale. The solution could be more than 1 individual instrument or module. | M
|  | Sampling | Ability to take single tubes or multiple samples preferably in a variety of formats e.g. 96, 48 and 24 well plates | M
|  | Sample volume | The smaller the volume required for data gathering the more desirable | M
|  | Sample processing | Samples should require no manipulation (apart from dilution in buffer) prior to running on the instruments. The shorter the time to results the more desirable. | M
|  | Key metabolites | Glucose, lactate, ammonia, glutamine, glutamate, Na+, K+, LDH | M
|  | Gases | pCO2, pO2 in culture medium. | M
|  | Other process variables | pH, osmolality | D
|  | Integration ability | Able to link to automated sampling systems | D
|  | Data format | Able to batch export results in csv or equivalent | D
|  | Connectivity | Able to link with LMS or other knowledge management system | D
|  | Consumables | Preferably able to use generic Eppendorf tubes, microtitre plates, etc | D
|  |  | Reagent costs for 500 samples must be estimated with typical lead times. | M
|  | Support | Evidence of good engineer backup and service contract availability. A 3 year maintenance agreement should be costed | M
|  | Training | Training to include calibration, operation, routine maintenance and troubleshooting | M
|  | Spares | Critical spares to be included | M
Installation & Utilities
|  | Installation | Costs for installation at NBMC, Darlington should be included | M
|  | Site requirements | Size requirements should be detailed as should utilies required for function. | M
|  | Electrical Power | Requirements should be detailed | M
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com

## Lot 6 - Cell Counters
User Requirement Specification |  | CPI Internal Tender Reference & Title:- | 1088 | Lot 6 – Automated Cell Counters
The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 40%
USER REQUIREMENT |  |  |  | Tenderer Response
Requirement codes |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
| System | Benchtop automated cell counters for the quantification of mammalian cells using a trypan blue based system | M
| Sampling | Ability to take single tubes or multiple samples | M
| Sample volume | The smaller the volume required the more desirable | M
| Sample processing | Samples should require no manipulation (apart from dilution in buffer) prior to running on the instrument | M
| Key parameters | Viable cell count, total cell count, % viable cells, cell diameter, aggregation/clumping, morphology (e.g. circularity) | M
| Output display | Simple display of parameters measured on screen. Plot of time course would be desirable. Data must be backed up, exportable and available for reanalysis | M
| Connectivity | Must be able to link with ambrTM 250 and the LMS or equivalent data management system | M
| Consumables | Preferably able to use generic Eppendorf tubes or similar. Reagent costs for 500 samples must be estimated and costed | M
| Quality control | Must be some means of monitoring system performance over time | M
| Support | Evidence of good engineer backup and service contract availability. A 3 year maintenance agreement should be costed | M
| Training | Training to include calibration, operation, routine maintenance and troubleshooting. Must also include software training | M
| Spares | Critical spares to be included | M
Installation & Utilities
| Installation | Costs for installation at NBMC, Darlington should be included | M
| Site requirements | Size requirements should be detailed as should utilies required for function. | M
| Electrical Power | Requirements should be detailed | M
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com

## Lot 7- Flow Cytometer
User Requirement Specification |  | CPI Internal Tender Reference & Title:- | 1088 | Lot 7 – Flow Cytometers
The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 40%
USER REQUIREMENT |  |  |  | Tenderer Response
Requirement codes |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
| System | A flow cytometry based instrument. Ability to interrogate single cells in a stream to build a picture of the heterogeneity of populations. Should be able to quantify mammalian and microbial cells | M
| Sampling | Ability to take single tubes or multiple samples preferably in a variety of formats e.g. 96, 48 and 24 well plates | M
| Sample volume | The smaller the volume required for interrogation the more desirable | M
| Lasers/LEDS | Up to 3 solid-state lasers or LEDS, one of which must be blue (excitation at 488nm) there must be flexibility in laser/LED configurations to allow maximum flexibility | M
| Detectors | Sufficient number of channels and detectors for the simultaneous measurement of at least 5 parameters by light scatter or fluorescent emission | M
| Functional assays | Apoptosis, cell viability, GFP expression, cell cycle, protein expression, mitochondrial membrane potential. System must be flexible with provision for novel assays (see below) | M
| Novel assays | System should be flexible enough to be used for novel assays if cells are labelled with appropriate biomarkers. | M
| Software | Must be intuitive and allow data acquisition and analysis. There should be scope for automating some data acquisition. | M
|  | Licence costs for software must be included. Data should be easily exported and backed-up. | M
| Connectivity | Able to link with LMS or other knowledge management system | D
| Consumables | Preferably able to use generic tubes, microtitre plates or similar | D
|  | Consumable costs for 300 samples must be estimated | M
| Support | Evidence of good engineer backup and service contract availability. A 3 year maintenance agreement should be costed | M
| Training | Training to include operation (including software), routine maintenance and troubleshooting | M
| Spares | Critical spares to be included | M
Installation & Utilities
| Installation | Costs for installation at NBMC, Darlington should be included | M
| Site requirements | Size requirements should be detailed as should utilies required for function. | M
| Electrical Power | Requirements should be detailed | M
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com

## Lot 9 (hidden)
User Requirement Specification |  | CPI Internal Tender Reference & Title:- |  | Lot 1 - Microscale vessels for mammalian culture
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 40%
USER REQUIREMENT |  |  |  | Tenderer Response
Requirement codes |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
| System | A flow cytometry based instrument. Ability to interrogate single cells in a stream to build a picture of the heterogeneity of populations. Should be able to quantify mammalian and microbial cells | M
| Sampling | Ability to take single tubes or multiple samples preferably in a variety of formats e.g. 96, 48 and 24 well plates | M
| Sample volume | The smaller the volume required for interrogation the more desirable | M
| Lasers/LEDS | Up to 3 solid-state lasers or LEDS, one of which must be blue (excitation at 488nm) there must be flexibility in laser/LED configurations to allow maximum flexibility | M
| Detectors | Sufficient number of channels and detectors for the simultaneous measurement of at least 5 parameters by light scatter or fluorescent emission | M
| Functional assays | Apoptosis, cell viability, GFP expression, cell cycle, protein expression, mitochondrial membrane potential. System must be flexible with provision for novel assays (see below) | M
| Novel assays | System should be flexible enough to be used for novel assays if cells are labelled with appropriate biomarkers. | M
| Software | Must be intuitive and allow data acquisition and analysis. There should be scope for automating some data acquisition. | M
|  | Licence costs for software must be included. Data should be easily exported and backed-up. | M
| Connectivity | Able to link with LMS or other knowledge management system | M
| Consumables | Preferably able to use generic tubes, microtitre plates or similar | M
|  | Consumable costs for 500 samples must be estimated | O
| Support | Evidence of good engineer backup and service contract availability | M
| Training | Training to include operation (including software), routine maintenance and troubleshooting | M
| Spares | Critical spares to be included | M
Installation & Utilities
| Installation | Costs for installation at NBMC, Darlington should be included | M
| Site requirements | Size requirements should be detailed as should utilies required for function. | M
| Electrical Power | Requirements should be detailed | M
OFFICE USE ONLY - This section below this line is for CPI Evaluation and Internal USE only
| Notes for Evaluation Team: | Name of Bidder to be inserted here | 0 | Lot 1 - Microscale vessels for mammalian culture
Please assess all of the answers given, with particular attention to the Mandatory answers provided, taking into account, the Desirable and Optional answers too, however, you do not need to score each individual answer, but based on the scoring range 0-5, just give one singluar score (default set at "5" in the red and yellow box below, and as we have to provide feedback to bidders, you must also provide comments on which section(s) above where you felt that they did not meet the requirements.
| Evaluation Panel | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com