# SCHEDULE 2 – THE SERVICES

- Service Specifications

Mandatory headings 1 – 4: mandatory but detail for local determination and agreement

Optional headings 5-7: optional to use, detail for local determination and agreement.

All subheadings for local determination and agreement

| Service Specification No. | Draft version 0.3 |
| --- | --- |
| Service | Wheelchair Repair & Maintenance for Herefordshire Clinical Commissioning Group |
| Commissioner Lead | Adrian Griffiths Head of Commercial Development Herefordshire Clinical Commissioning Group |
| Provider Lead |  |
| Period | 1st April 2018 – 31st March 2023 |
| Date of Review | At the request of either Commissioner or Provider |

SECTION 1. Population Needs

- National/local context and evidence base The number of people with disabilities is rising, it is estimated that there are 1.2million wheelchair users in England – just over 2% of the population. Wye Valley NHS Trust operates a Wheelchair Service to provide wheelchair services to children and adults who are registered with a Herefordshire GP (“Service Users”). This includes Service Users whose normal place of residence is within the geographical boundaries of the county of Herefordshire and those in neighbouring counties. Temporary Resident Service Users are also entitled to access Repair Service only. Herefordshire Clinical Commissioning Group requires a Contractor to deliver the Service to Service Users, at one of the Service Delivery Locations detailed in Appendix 2 below, covering one or a combination of the following Services: - delivery & collection service – Section 2 - storage & stockholding service – Section 3 - repair service – Section 4 - reconditioning service – Section 5 - planned preventative maintenance service powered wheelchairs & manual wheelchairs – Section 6 For the avoidance of doubt, the purchase of all new wheelchairs is excluded from the Contract.

2. Outcomes

2.1 NHS Outcomes Framework Domains & Indicators Domain 1; Preventing people from dying prematurely. Domain 2; Enhancing quality of life for people with long-term conditions; √. Domain 3; Helping people to recover from episodes of ill-health or following injury; √. Domain 4; Ensuring people have a positive experience of care; √. Domain 5; Treating and caring for people in safe environment and protecting them from avoidable harm; √. 2.2 Local defined outcomes The Contractor must deliver the Contract in accordance with the Medical Devices Agency Controls Assurance Standard April 2008 (MDA-CAS), which specifies individual criteria in respect of each of the Sections of this Contract.

3. Scope

3.1 Aims and objectives of service To provide a repair and planned preventative maintenance service for the population of Herefordshire who are users of NHS supplied wheelchairs (in line with Herefordshire Clinical Commissioning Group criteria). 3.2 Service description/care pathway Service Availability The Contractor will ensure the Service is available and fully operational between Monday to Friday 08.30 to 17.00 including Public Holidays. In addition, the Contractor will provide an emergency repair service only (Section 4) from 08.00 to 23.00, 7 days a week, including Public Holidays. This constitutes a Key Performance Indicator. Equipment Covered by the Service to be Provided by the Contractor "The Equipment" is defined as the wheelchairs detailed in this paragraph plus any accessories. There are a range of companies by whom the equipment is made and the Contractor must provide the Service in respect of any Authority owned Equipment regardless of the manufacturer. Assessment Wheelchairs These are wheelchairs owned by Herefordshire Clinical Commissioning Group and used solely by Herefordshire Clinical Commissioning Group for the assessment of Service Users. The wheelchairs can be manual, powered or buggy. Asset Wheelchairs These are Authority owned wheelchairs located in community hospitals throughout Herefordshire for in-patient use. Currently, Herefordshire Clinical Commissioning Group owns approx. 100 such wheelchairs which are manual, self-propelling or transit wheelchairs, the majority of which are of Remploy design. The Contractor will be required to provide the Service within the community hospital premises. This list is not exhaustive and is subject to change throughout the Contract Period. Referral & Registration Process The Wheelchair Service is operated by Wye Valley NHS Trust and is based at Ruckhall Lane, Belmont, Hereford, HR2 9RP. The Wheelchair Service’s primary role is to undertake clinical assessments of Service Users in accordance with the devolution of the Service to the NHS in 1991, to supply Equipment to these Service Users, and to manage this Contract which supports these obligations. The Wheelchair Service receives a referral from healthcare professionals, and then: - registers the Service User's details onto the Wheelchair Service's database and informs the Contractor; and - undertakes a clinical assessment of need which includes either a domiciliary visit or involves the Service User attending a specialist clinic operated by the Wheelchair Service. Where a Service need has been identified, the Wheelchair Service will raise a Works Order to the Contractor containing instructions to provide a Service. The Works Order will contain as a minimum: - Service User name, address and contact details and the Service Delivery Location if not the Service User’s normal place of residence - equipment number and type to be provided - Service to be provided - special delivery or installation instructions e.g. English is not the first language, deaf, etc. - if an emergency service is required. On receipt of Works Order the Contractor must: - provide the Service as detailed in the Works Order in accordance with the relevant Section(s) - complete the Service to the Service User's satisfaction - immediately forward the completed Works Order to the Wheelchair Service (works orders must be made available to Herefordshire Service Provider if required for checking) - use the information contained within the Works Order to provide accurate management information defined in Section 15 below. In the case of a part complete Works Order, e.g., awaiting non stock spares, the Contractor must retain the Works Order until it is completed, then follow the procedure above. Reporting Process for Sections 2, 4 & 7 (Delivery & Collections, Repairs & PPM) The Contractor must provide and maintain a telephone contact number available during the times detailed above to enable Service Users to directly report the need for Service to the Contractor. Only Service Users who are registered with the Wheelchair Service are eligible to directly contact the Contractor and receive these Services. No separate additional Works Orders will be issued by the Wheelchair Service and the Contractor must ensure all Service provided is accurately logged for reporting under Section 16 below. The Contractor must respond to the Service User within one hour following receipt of the call, and the Service must be completed as specified in Sections 2, 4, & 7. Response Times These are detailed as Key Performance Indicators in Sections 2 to 8 inclusive. Work In Progress Prior to 1 April 2018 For the avoidance of doubt, any Service not completed as at 31 March 2017 will be the responsibility of the current contractor to deliver and complete. Transition Period Between date of Contract award and the Contract Commencement Date a period has been allocated to enable the successful bidder to adequately prepare for the start of Service delivery and work with Herefordshire Clinical Commissioning Group to agree and sign the Contract document. As a minimum, the successful bidder will: - undertake the agreed Transition Plan; - implement any new ICT proposal (see Section 8) where adopted by Herefordshire Clinical Commissioning Group; - meet with the Herefordshire Wheelchair User Group; and - produce key contract details/literature to explain the Service to be delivered including the contact telephone number, the Contractor's general contact number, details of its complaints procedure, and the Wheelchair Service's contact details. This will be produced and agreed with Herefordshire Clinical Commissioning Group's Contract Manager by the end of February 2018 and issued in March 2018 to Service Users. SECTION 2 - DELIVERY & COLLECTION SERVICE Delivery and Collection This includes collection of Wheelchair Service owned Equipment from Service Users to the Contractor's storage facility (see Section 3 below) and delivery to individual Service Users from the Contractor's storage facility and to and from Herefordshire Wheelchair Services premises. The Contractor will prepare the Equipment prior to delivery in accordance with the manufacturer's instructions and/or Wheelchair Service guidelines. Preparation shall include, but not be limited to: - all necessary adjustments - interchanging of appropriate parts - fitting of accessories - fitting/recording all serial or identification numbers on the Works Order - supplying and fitting Wheelchair Service provided identification labels - Issuing user manual. The Wheelchair Service will provide the Contractor with a bulk supply of the necessary identification labels. The Contractor must notify the Wheelchair Service 5 days in advance and in writing when a replacement supply is needed. Collected Equipment The Contractor will: - collect the Equipment from the Service User at the address detailed in the Works Order and take for storage to the Contractor's facility - compile a list of all Equipment collected from Service Users and forward weekly by 5pm on Mondays to the Wheelchair Service detailing the information necessary to uniquely identify each wheelchair or accessory; - enable access by the Wheelchair Service to the storage facility on a weekly basis (dates and times to be jointly agreed between the parties) to enable the Wheelchair Service to inspect the Equipment, categorise it, and instruct the Contractor to: - decontaminate the Equipment as specified in Section 5 and either retain in the storage facility or deliver to a Service User - recondition the Equipment as specified in Section 6 and either retain in the storage facility or deliver to a Service User - scrap the Equipment. The Contractor will be responsible for the scrapping any designated Equipment in accordance with current legislation. Delivery & Handover of Equipment to Service Users – to include new and re-conditioned equipment. Pre Delivery of Equipment to a Service User The Contractor must contact the Service User to arrange a mutually convenient time to deliver and handover the Equipment and meet the Key Performance Indicators detailed below. The Equipment must be in working order and as specified in the Works Order. In addition, the Contractor must ensure that any powered Equipment and the associated battery chargers are subjected to a commissioning test in accordance with the manufacturer's instructions. Delivery of the Equipment At the time of delivery, all Equipment will be handed over to the Service User (and/or their representative) in accordance with the manufacturer's instructions and relevant handover guidelines as detailed below. The Equipment will then become Equipment to be inspected annually by the Contractor where relevant and in accordance with MDA-CAS 2008 and as specified in Section 7. Handover of Equipment - Powered Wheelchairs The handover procedure for powered wheelchairs by the Contractor is: - upon arrival, the Contractor's staff must establish their identity by showing ID cards and stating the reason for the visit - transfer the modifications or accessories from the old wheelchair if appropriate - demonstrate to the Service User the use of the wheelchair as follows: - method of battery access - battery maintenance procedure - battery charging procedure - leave for retention by the Service User a copy of the - manufacturer's operating manual, and charging and battery - maintenance instructions. - demonstrate other aspects of the Equipment which the Service user would normally need to use: - adjustment of armrest height where adjustment is provided - the folding action and release mechanism of the swinging detachable footrests - adjustment of footrest height - any seating adjustment such as height which may be fitted - operation of the parking brakes - engagement and disengagement of the motors. - with the Service User seated in the wheelchair (and any representative present), demonstrate or adjust the following: - adjust the height of the footrests and other items if required - layout and operation of driving methods - advise the Wheelchair Service immediately if the control box needs to be re-programmed. - report immediately any environmental issues, hazards or unsatisfactory Equipment to the Service User, any representative present, and the Wheelchair Service. - Completion of handover certificate as per manual chair – because if contractor is carrying out this activity there should be written confirmation which is more important with power chairs. The Service User must drive the wheelchair under power in the presence of the Contractor's staff in order for any further adjustments to be made. Handover of Equipment - Manual Wheelchairs The handover procedure for manual wheelchairs by the Contractor is: - upon arrival, the Contractor's staff must establish their identity by showing ID cards and stating the reason for the visit - demonstrate to the Service User the use of the wheelchair as follows: - how to enter and leave the wheelchair - the folding mechanism, how to fold and unfold the wheelchair - the folding backrest, release of locking pins - removal and replacement of the detachable armrests, where fitted - the release mechanism of the swinging detachable footrests - adjustment of the footrest height - application of the brakes - how to use the pump (where supplied) - fit accessories, if provided and adjust to suit client - with the Service User seated in the wheelchair (and any representative present) adjust the height of the footrests to suit - leave for retention by the Service User a copy of the manufacturer's - operating manual and any maintenance instructions - report immediately any environmental issues, hazards or - unsatisfactory Equipment to the Service User, any representative - present, and the Wheelchair Service - complete and return immediately to the Wheelchair Service the Handover Completion Certificate (see Appendix 1). Key Performance Indicators The Contractor must complete each Works Order and to the Service User's satisfaction: - within 5 working days following issue of the Works Order (NOTE: day 1 commences the first working day following issue of the Works Order); OR - within 24 hours following issue of the Works Order where 'emergency action' is required (as instructed in the Works Order). In the case of fast-tracking palliative (end of life) and to facilitate hospital discharges a quicker timescale of 48 hours is required. An 'emergency action' is defined where: - a Service User is dependent upon their wheelchair for mobility indoors; or - the condition of the wheelchair is such that the wheelchair is not in a mobile and safe condition. - clients living alone and at risk of harm If a loan chair of suitable standard is available then the emergency timescale can be reduced. SECTION 3 – STORAGE & STOCKHOLDING SERVICE Storage of Equipment The Contractor will provide storage for Herefordshire Clinical Commissioning Group owned Equipment (wheelchairs and accessories) where detailed in a Works Order and will include transfer of current contractor's stock - equipment and accessories. It is not a requirement of this Contract that the storage facility must be located in Herefordshire but it must be located no more than a maximum one hour's driving time of the Wheelchair Service's base in Hereford. This is to ensure that when equipment needs accessing, checking or input in any other way from the Herefordshire Wheelchair Service personnel do not have to spend too much time travelling, taking away from clinical services provision time. Equipment must be stored at all times in a secure and covered environment which offers facilities to ensure the maintenance of the Equipment in accordance with MDA-CAS. The Contractor must insure all the Wheelchair Service's Equipment held in storage to cover, as a minimum, fire, damage, theft and loss at an appropriate level to provide for New for Old replacements. The Contractor, upon request of Herefordshire Clinical Commissioning Group and/or the Wheelchair Service, will produce documentary evidence that such insurance is properly maintained at all times. Relinquished Equipment will not count towards the Wheelchair Services' stockholding until the Equipment has been re-conditioned and the Wheelchair Service notified of its readiness for inspection. For new Equipment purchased by Herefordshire Clinical Commissioning Group, the Contractor will forward to Herefordshire Clinical Commissioning Group (the Wheelchair Service) copies of delivery notes for Equipment received by the Contractor within three working days of receipt. The Contractor will be responsible for disposing of all packaging in an appropriate manner in accordance with MDA-CAS, Section 10 below, and relevant legislation. Within 48 hours of Equipment being collected, the Contractor must unpack, check and verify the condition of the Equipment, including setting aside for inspection by the Wheelchair Service and where appropriate, re-pack and storage. The Wheelchair Service must be notified in writing of any missing, damaged or incorrect Equipment immediately the fault is identified. Purchase of Equipment - Accessories & Spare Parts (Section 4) The Contractor is responsible for the purchase and supply of accessories and spare parts required to fulfil individual Works Orders and purchase of replacements equipment due to repair and meet the Key Performance Indicators detailed in each Section of this Specification. Accessories typically include: - Backrest extensions - Elevating leg rests - Stump boards - Detachable trays - Waist belts - Calf straps - Clip-on footboard extensions - Solid rear tyres - Clear spoke guards - Bexhill armrests - Standard Anti-tipping levers There is no requirement for the Contractor to maintain a stockholding of accessories in a storage facility. A Just In Time Supply system of operation is acceptable as long as the Key Performance Indicators specified for each Section are met. All accessories must be the manufacturers' brand or equivalent alternative. Stock Control The Contractor shall at all time maintain an up to date stock control system, which provides an audit trail for tracing the movement of Equipment booked into store, issued out, and stored. This includes all Authority owned Equipment (wheelchairs and accessories) and any accessories kept as stock by the Contractor. The records must be available for inspection by the Wheelchair Service upon demand at all times. Stock Checks The Wheelchair Service will undertake checks of their own stock, at the Contractor's premises, either on an unannounced spot check basis or by prior arrangement. Stock checks will also form part of the Quality Audit undertaken annually by the Wheelchair Service's representatives, and the Contractor must ensure access to the storage facility, inspection of Equipment and stock control system by these representatives. Following the stock check, discrepancies will be notified to the Contractor who must investigate and report its outcome to the Wheelchair Service in writing within 5 working days. In respect of any Equipment loss or damage, Herefordshire Clinical Commissioning Group will charge the Contractor for replacement Equipment, set at either the manufacturer's cost or the cost of the reconditioned Equipment, plus delivery and a 15% handling charge. SECTION 4 – REPAIR SERVICE The Contractor will be required to repair the Equipment defined above as detailed in a Works Order. Replacement Parts The Contractor will use replacement parts that match those specified by the manufacturer or are an equivalent alternative. Equipment Pre 1991 For guidance on Equipment manufactured pre April 1991, the Contractor must contact the Wheelchair Service. All repairs, maintenance and adjustments to Equipment manufactured after April 1991 will be carried out in accordance with the procedures set out in the relevant manufacturer's technical manual. Equipment Misuse The Contractor will report any cases of Equipment misuse, neglect or repeated failure immediately in writing to the Wheelchair Service. Salvaged Items The Contractor will only supply items of Equipment salvaged from previously scrapped equipment following a risk assessment (DB9801) and prior written agreement from the Wheelchair Service. Serviceable Stock Parts These will not be removed from any stock held by the Contractor unless written agreement is given by the Wheelchair Service. Special Seating The Contractor will be required to have availability of spare parts for special seating requirements. It is envisaged that in many cases, emergency repairs will need to be carried out. The Contractor will ensure that staff members are fully cognisant and competent in such repairs. Herefordshire Clinical Commissioning Group will make available to the Contractor, where requested free of charge training for special seating repairs. Repairs to Powered Equipment (outdoor and indoor/outdoor wheelchairs) In the event of a breakdown of a powered wheelchair away from the premises detailed in the Service Delivery Locations (in Appendix 2 below), a Service User will be responsible for making their own arrangements for getting themselves and their Equipment to their normal place of residence in accordance with their conditions of supply. The Equipment will then be repaired by the Contractor as specified in Section 2 above. Warranty Repairs Equipment purchased by Herefordshire Clinical Commissioning Group is generally covered by a 12 month warranty period or longer. If a repair is needed to the Equipment during the warranty period, the Contractor must contact the Wheelchair Service immediately. The Wheelchair Service will liaise with the manufacturer and advise the Contractor if the defective part needs to be returned by the Contractor to the manufacturer or any other action is required. Product Recalls Product recalls may occur within or outside of the manufacturer's warranty period. The Contractor must abide by the manufacturer's instructions pertaining to each individual event, and immediately inform the Wheelchair Service in writing. Voucher Scheme A Service User's clinical needs will be assessed by the Wheelchair Service, and a type Equipment recommended. If the Service User wishes to have a different choice of Equipment type, and the preferred Equipment exceeds the cost of the recommended Equipment, then a voucher scheme is available to the Service User. This entitles the Service User to directly purchase the preferred Equipment with the Wheelchair Service contributing up to the value of the recommended Equipment. The Equipment then becomes the property of the Service User and is designated as privately owned Equipment which is excluded from all sections of this Contract. Note that a Service User may have access to both recommended and preferred Equipment. Batteries, Battery Charges and Control Boxes These can be either the manufacturer's recommended brand or an equivalent alternative. These must be available, whether as stock or non stock items, to enable the Contractor to complete the Key Performance Indicators detailed below. Key Performance Indicators The Contractor must complete each Works Order and to the Service User's satisfaction: - within 3 working days following issue of the Works Order (NOTE: day 1 commences the first working day following issue of the Works Order); OR - within 24 hours following issue of the Works Order where 'emergency action' is required (as instructed in the Works Order) An 'emergency action' is defined where: - a Service User is dependent upon their wheelchair for mobility indoors; or - the condition of the wheelchair is such that the wheelchair is not in a mobile and safe condition. SECTION 5 – DECONTAMINATION SERVICE – includes all equipment, cushions and accessories for Herefordshire Wheelchair Service. Decontamination Procedure Where instructed in a Works Order, the Contractor will collect the Equipment from the Service User and undertake a decontamination procedure. The Contractor must ensure a decontamination procedure is in place at all times which complies with the Health Service Guideline HSG (93) 26 and Controls Assurance Standard – Decontamination of Medical Devices – Criterion 7. Key Performance Indicators The Contractor must complete each Works Order as follows: - collect the Equipment within 5 working days following issue of the Works Order (NOTE: day 1 commences the first working day following issue of the Works Order); and - complete the decontamination procedure within 1 working day following collection. SECTION 6 – RE-CONDITIONING SERVICE Re-conditioning of Equipment means bringing a wheelchair (manual or powered) back to a condition that is suitable and safe for issue to any Service User. The Wheelchair Service will assess whether Equipment is suitable for re-conditioning. Procedure The Wheelchair Service will advise the Contractor in writing of the specification of the re-conditioning required, and a quotation must be returned to the Wheelchair Service in 3 working days following receipt of the specification. The Wheelchair Service will consider the quotation and inform the Contractor, via issue of a Works Order, if the work is to be undertaken. The Wheelchair Service is not obliged to accept the Contractor's quotation and may use an alternative provider(s) if it chooses (and/or can demonstrate greater value for money in doing so) or will instruct the Contractor to scrap the wheelchair as specified in Section 2 above. Parts Where the Contractor is instructed in the Works Order to provide a Service, the parts used to re-condition the Equipment can either be the manufacturer's recommended brand or an equivalent alternative. Warranty Equipment re-conditioned by the Contractor will be subject to a 6 month warranty from date of delivery to a Service User against faulty workmanship and replacement parts and/or materials. Such defects must be immediately rectified, free of charge, by the Contractor. Inspection The Wheelchair Service will inspect all completed re-conditioned wheelchairs prior to either return to stock and storage or delivery to a Service User. The Contractor must immediately contact the Wheelchair Service when re-conditioning has been completed to arrange for inspection to take place. Labels The Contractor will affix to the re-conditioned wheelchair a label identifying the Contractor's company by name and the date of re-conditioning (separate to Herefordshire Wheelchair Service ID labels). Key Performance Indicators The Contractor must complete each Works Order as follows: - submit the quotation to the Wheelchair Service within 3 working days; following receipt of the specification (NOTE: day 1 commences the first working day following issue of the Works Order); and - complete the wheelchair re-conditioning within 10 working day following receipt of a Works Order ready for inspection. SECTION 7 – PLANNED PREVENTATIVE MAINTENANCE SERVICE (PPMS) Eligible Wheelchairs For each year of the Contract, the Contractor will be responsible for determining which powered Equipment that has been issued Service Users qualifies for annual PPMS, and undertake this Service on or as near to the anniversary date of the first annual PPMS. PPMS Requirements In respect of each powered wheelchair, the Contractor will undertake, as a minimum, the following: - service the wheelchair as per the manufacturer's recommendation; - replace any part that is not fit for purpose; - act at all times in accordance with the MDA-CAS; - undertake the PPMS either at a Service User's normal place of residence or at the Contractor's premises. Where the latter takes place, the Contractor will be responsible for collection and delivery of the wheelchair. Parts The parts used to undertake the PPMS can either be the manufacturer's recommended brand or an equivalent alternative. Annual PPMS Programme The Contractor shall submit their proposed annual PPMS programme to the Wheelchair Service by 31 March each year. The programme must enable each Equipment qualifying for Service to be uniquely identified, the date schedule for the Service to undertaken, and Service User details. Review The Contractor's progress in delivering the annual PPMS Programme will be reviewed quarterly as specified in Section 15 below. Annual PPMS Programme for Manual Wheelchairs The Contractor may be required to undertake an annual PPMS on manual wheelchairs in addition to powered wheelchairs. The Contractor will be advised at time of Contract award whether this Service is required and included in the Contract. If included, the Service will be undertaken as specified in this Section 8. SECTION 8 – ICT The Wheelchair Service currently uses, a Sam Shaw Equipment & Patient Management System ("the System") to operate the Wheelchair Service. In summary, the system provides: - a database holding Service Users' personal information - for the issue of Works Orders - for the management of Service User clinics - for the creation and issue of letters - for the creation of statistical and bespoke reports - for the extraction of statistical data, equipment logs and asset identification. The current Specification is: - Office 2000/2003 - Workstation One® Intel™ Core i5 4GB RAM - 250GB Hard Disk, - Wide Screen 18” Monitor - Windows 7 Professional 64bit - 100MB Network Card - Server SQL Server 2008 R2 Std Edition 8 GB RAM - Network 100MB CAT 5 UTP - 100MB Hub The Wheelchair Service owns the data contained within the System and Sam Shaw owns the operating system. This System and the current contractor's management information system operate independently and there is no interaction. The Contractor is responsible for providing and maintaining its own management information, stock control and quality assurance systems in order to meet the Service and contract management requirements (please see Section 15 below). The Contractor's systems must hold records in compliance with the ISO15489 Records Management standard. Where viable, the Contractor's systems should be compatible with and interact with Herefordshire Clinical Commissioning Group's System. Alternative Equipment & Patient Management System Herefordshire Clinical Commissioning Group wishes to consider the viability of adopting an alternative Equipment and Patient Management System, preferably web based which interacts with the Contractor's management information, stock control and quality assurance systems to produce a system that is seamless, operational 24/7, and accessible at all times to Herefordshire Clinical Commissioning Group's Wheelchair Service. The system must also enable electronic transactions including the transmission of Works Orders to the Contractor. The proposed solution does not have to run in the NHS Herefordshire/Council environment but could be hosted elsewhere. NOTE: The proposal of an alternative system by bidders is optional but any proposal offered must be included in their response to the future Invitation Tender and the cost will be requested in the Pricing Schedule. SECTION 9 – INFORMATION GOVERNANCE The Contractor must comply at all times with the relevant legislation that affects the way data is held, processed, and released. The key legislative requirements are detailed below. Note that this list is not exhaustive and any amendment, update or new Act coming into force shall equally apply to the Contractor: - Freedom of Information Act 2000 - Data Protection Act 1998 - Environmental Information Regulations 2004 - E-Government regulations - Civil Contingency Act 2004 - Equality Act 2010 Personal Data The Contractor will notify Herefordshire Clinical Commissioning Group within five working days if it receives: - a request from a Data Subject to have access to that person's Personal Data; or - a complaint or request relating to Herefordshire Clinical Commissioning Group’s obligations under the Data Protection Legislation; and - transfer to Herefordshire Clinical Commissioning Group a copy of all information in the Contractor’s possession relating to the requested information. In order that Herefordshire Clinical Commissioning Group can respond to the request/s, the Contractor shall: - transfer to Herefordshire Clinical Commissioning Group all requests for Information that it receives within two working days of receiving a Request for Information in relation to the Freedom of Information Act 2000 and The Environmental Information Regulations 2004 and - provide Herefordshire Clinical Commissioning Group with a copy of all Information in the Contractor’s possession, within five working days requested by Herefordshire Clinical Commissioning Group in order to respond to a Freedom of Information Act 2000 and The Environmental Information Regulations 2004 request. SECTION 10 – CORPORATE SOCIAL RESPONSIBILITY Sustainability The Contractor will comply at all times with environmental legislation and must operate an environmental/sustainable policy that, as a minimum, demonstrates how the Contractor mitigates, minimises, and monitors the environmental impact of its business operations from pollutants, hazardous substances, and greenhouse gas emissions from energy and fuel use. The Contractor should possess EMS certification. The Contractor must ensure its staff are trained in sustainability. The Contractor will comply at all times with the packaging (Essential Requirements) Regulations 1998 (SI 1165) (as amended) which implements the EU Directive on Packaging and Packaging Waste (94/62/EC) in the UK, and which requires packaging to be minimised, recoverable, and not to exceed by weight specified concentrations of heavy metals. Equality & Diversity The Contractor will comply at all times with the Equality Act 2010. SECTION 11 – CONTACT & COMMUNICATION WITH SERVICE USERS Confidentiality and Information Governance to NHS Standards should be adhered to at all times. Further advice on these standards and training can be sought from Herefordshire Clinical Commissioning Group. Prior to delivering Service the Contractor must always contact the Service User and agree a mutually convenient time to undertake the Service. At the first visit only, the Contractor must leave with the Service User a copy of the key contract details provided. The Contractor is responsible for providing any updates to the key contract details as and when a change occurs, e.g., change to the contact telephone number. All such changes must be agreed in writing with the Wheelchair Service prior to the Contractor issuing to Service Users. The Contractor must at all times ensure their staff behave in a courteous manner when addressing Service Users and/or their representatives. The Contractor's staff must always produce their Identity Card to a Service User immediately upon arrival at the Service Delivery Location. If upon arrival at the Service Delivery Location there is no response, the Contractor's staff must leave a card identifying the Contractor by name, and a telephone number requesting the Service User to call to re-arrange the appointment. The Contractor must maintain a complaints and disputes resolution procedure, and ensure its existence is communicated to Service Users. SECTION 12 – QUALITY ASSURANCE & CONTINUOUS IMPROVEMENT Quality Assurance The Contractor must operate and maintain at all times an approved third party assessed quality assurance scheme that meets the requirements of BS EN ISO150 900L 2008 (as updated from time to time). Herefordshire Clinical Commissioning Group's Wheelchair Service will undertake an annual audit of the scheme and the performance of the Contract. Service User & Carer Feedback The Contractor must seek views, retain and action. Further details to be provided. Continuous Improvement The Contractor must implement a programme of continuous improvement, including benchmarking, to operate during the Contract period. The programme must include, but not be limited to: - a description of the system - a project plan for introducing and developing the system - realisable financial efficiencies and Service improvements that benefits Service Users which will be delivered annually. SECTION 13 – STAFFING Contract Manager The Contractor will allocate an individual as the designated Contract Manager to manage the Contract and act as Herefordshire Clinical Commissioning Group’s central point of contact. The individual is not required to be dedicated full-time to the management of this Contract, but sufficient time must be allocated to ensure the Contract is delivered in the manner and to the standards contained in this Specification. The Contractor will be responsible for allocating a sufficient number of qualified and technically proficient staff to perform the Contract, especially with regards to the engineers who must be familiar with the diagnostic tools associated with powered wheelchairs. Staff Recruitment & Security Checks Where a new staff member is recruited, the Contractor must ensure the requirements of all relevant legislation and best practice are adhered to. All staff delivering a Service must have Criminal Records Bureau Enhanced clearance prior to delivering Service under the Contract. Identity Cards The Contractor shall provide and issue to all staff who liaise with Service Users plastic sealed identity cards to a standard approved by the Wheelchair Service. As a minimum, each identity card must include a current photographic likeness of each staff member together with the Contractor’s name, address and telephone number. The identify card must always be shown for inspection when visiting a Service User. The Contractor must provide upon request to the Wheelchair Service a listing of authorised staff members with identity cards delivering Service under the Contract. Where staff changes occur, the Contractor shall immediately update the listing of authorised staff and inform the Wheelchair Service. Staff Training & Continuous Professional Development Training The Contractor must ensure appropriate annual staff training takes place and in accordance with MDA-CAS. With regards to the engineers, this must include (but is not limited to) ensuring that they undergo training and continuous professional development which includes awareness of developments in industry and best practice. Safeguarding Prior to commencing delivery of Service, the Contractor must ensure all staff are trained in Herefordshire Clinical Commissioning Group’s Safeguarding Policy including any revised or updated Policy issued by Herefordshire Clinical Commissioning Group. Absence Cover In the absence of any staff member allocated to deliver the Contract, for whatever reason, the Contractor must immediately provide a suitably qualified alternative staff member who at all times meets the requirements of this Section 13. TUPE Bidders should note that TUPE could to apply to this Contract. The application of TUPE is for the current contractor and the new Contractor to determine, but on a voluntary basis, the current contractor. SECTION 14 – BUSINESS CONTINUITY & EXIT PLANNING Business Continuity Plan The Contractor must ensure at all times that an up to date and robust Business Continuity Plan is in place. The Plan will be reviewed by the Wheelchair Service (together with other Authority representatives where appropriate) at the Annual Contract Review. Exit Plan The Contractor must ensure at all times that an up to date and robust Exit Plan is in place that also complies with (but are not limited to) clauses 27 and 28 of the NHS Conditions of Contract for the Supply of Services. The Exit Plan will operate either if the Contract is terminated early for whatever reason or at the end of the Contract Period (including any extension period(s)). The Plan will be reviewed with the Wheelchair Service (together with other Authority representatives where appropriate) at the Annual Contract Review. SECTION 15 – CONTRACT MANAGEMENT Contract Management Herefordshire Clinical Commissioning Group (via a commissioning representative and the Wheelchair Service) is responsible for the management of the Contract which will be undertaken through the use of various performance measures in conjunction with the parties detailed below. The Contractor will be expected to work proactively to implement effective contract management. Roles & Responsibilities Key roles and responsibilities of the various parties, activities required to be undertaken, and specific performance and reporting measures are detailed below. Key Roles & Responsibilities Herefordshire Clinical Commissioning Group undertake contract management including: - monitor Contractor performance against the Service and Key Performance Indicators together with the Herefordshire Wheelchair Service User Group representatives - co-ordination of management information from the Contractor - managing performance reviews with the Contractor - managing any performance issues - managing Contract variations and disseminating outcomes to Service - Users and internally within Herefordshire Clinical Commissioning - Group - managing the extension of any optional Contract extension period(s) - facilitating and promoting continuous improvement the Contractor shall contribute to the contract management process by: - delivering contractual obligations and individual Works Orders - providing management information and in a timely fashion to Herefordshire Clinical Commissioning Group (and as requested on an ad hoc basis) - attending performance review meetings - providing an effective account management and customer services role to both Herefordshire Clinical Commissioning Group and Service Users - contributing to continuous improvement and innovation leading to improved value for money and process efficiencies via its Continuous Improvement Programme. Quarterly & Annual Review Meetings The Contractor must attend quarterly performance review meetings and an annual Contract review meeting with Herefordshire Clinical Commissioning Group. The meetings will be held on Herefordshire Clinical Commissioning Group's premises and, as a minimum, will be attended by the Contractor's Contract Manager. Submission of Performance Reports Quarterly Reporting Following commencement of the Contract, the Contractor will provide a performance report to Herefordshire Clinical Commissioning Group, and to the Wheelchair Service, on a quarterly basis to include the following information/data: - Number of repairs broken down by manual and powered electric wheelchairs, number outstanding, number direct from Service Users and Temp Service Users - Number of product recalls - Number of collections & deliveries - Number of PPMs undertaken, number programmed, number completed, service user feedback - Equipment (wheelchairs and accessories) stock holding and any spare parts - Number of emergency call outs for repairs. Delivery and collection as per Section 2 - Number of completed powered and manual wheelchairs re-conditioned - An Exception Report - Number of chairs scrapped listed by manufacturer, model number and serial number - Number of warranty repair claims reported, action taken and outcome - Number of Service Availability failures, with reasons - Number of complaints received, action taken and outcome - Performance against all KPIs and individual KPI performance against the target with supporting breakdown - Staff training – sustainability, engineers, equality, safeguarding - Continuous improvement programme progress - Number occasions when service not available including reasons for this. - IG - ICT - Diversity & equality - Finance - Service user & carer feedback – service delivery and improvements Annual Reporting 1 Audit Report of the Quality System 2 Staff Training & turnover 3 CPD evidence 4 Business Continuity Plan 5 Exit Plan 6 Continuous Improvement Programme Reports must be submitted to Herefordshire Clinical Commissioning Group's nominated contact and to the Herefordshire Wheelchair Services nominated contact(s) within 10 working days after the end of each quarterly (or annual) reporting period. Inspection and Review of the Contractor’s Records Herefordshire Clinical Commissioning Group may undertake, at a frequency to be advised, an inspection and review of the Contractor’s records and performance reports to verify their contents. For the avoidance of doubt, this requirement will be in addition to any audit requirement detailed in the NHS Conditions of Contract for the Supply of Services. Annual Audit Herefordshire Clinical Commissioning Group or their representative will undertake an annual audit of the Contractor's quality assurance system to ensure it meets the requirements of BS EN ISO150 900L 2008 Quality System. Key Performance Indicators The Contract will be monitored against the Contract summarised below, and the Key Performance Indicators specified in detail in the table below: • communication and customer liaison; • service delivery; • service quality; • production of management information; • customer satisfaction; • corporate social responsibility; • IG • ICT • quality assurance • continuous improvement It is intended that this list will be developed throughout the lifetime of the Contract as demand evolves and will reviewed annually (as a minimum) in consultation with the Contractor. Summary of KPIs KPI 1 Section 1. - Category: Service Availability - Measure: To ensure Service Users have access to Service during prescribed times including public holidays - Target: 95% - Data Source: No. times Service unavailable KPI 2 Section 2. - Category: Service Delivery - Measure: To ensure Service Users receive a prompt service and an emergency service where need is identified by the Wheelchair Service - Target: 95% - Data Source: No. of times completion of Service missed KPI 3 Section 4. - Category: Service Delivery - Measure: To ensure Service Users receive a prompt Repair Service - Target: 95% - Data Source: No. of times completion of Service missed KPI 4 Section 6. - Category: Service Availability - Measure: To ensure the availability of Equipment to Service Users promptly following decontamination 95% - Data Source: No. of times completion of Service missed KPI 5 Section 1. - Data Source: MDA-CAS breaches KPI 6 Section 11. - Category: Customer Complaints - Measure: To ensure Service User feedback is addressed - Target: 95% - Data Source: Number and response to complaints received KPI Failure In respect of all KPIs, the Contractor must meet a minimum of 95% of the actual activity that occurred during the quarterly period. The consequence of not meeting 95% in respect of one or more KPI for a quarter period will be a 2.5% deduction from the quarterly block payment ex VAT to the Contractor by Herefordshire Clinical Commissioning Group. In addition, where applicable performance and breach clauses in the standard Ts&Cs may also be invoked. Appendix 1 HANDOVER CERTIFICATE FOR MANUAL WHEELCHAIRS The Contractor must complete and return to the Wheelchair Service this signed certificate immediately following completion of the Service to the Service User. Service User's Name: _____________ Ref No: _______________ Model:__________________________ Serial No: ________________ Please indicate in the box with a tick whether the task was completed i) Seat Service User in the wheelchair and adjust the footrest height ii) How to collapse and fold the wheelchair(s. iii) How to fold the backrest iv) Demonstrate the footrest mechanism v) How to release and replace the armrests vi) How to enter and leave the wheelchair vii) How to use the pump (where supplied) viii) How to operate the brakes (ix) Deliver the manufacturer's handbook to the Service User Signature of Service User: _____________________________________ Signature of Contractor: _____________________________________ Date: ________________________________________________________ Appendix 2 Population covered Geographical Coverage (source: Google Maps) Service Delivery Locations The list detailed below is not expected to change during the Contract Period, but any additional premises requiring Service (or no longer requiring a Service) will be notified to the Contractor. Service Users' homes and place of work Schools – special and mainstream Westfield School, Leominster Trinity School, Hereford Barrs Court School, Hereford Blackmarston School, Hereford Bishops Bluecoat School, Hereford All Wye Valley NHS Trust premises Day Centres Most Service Users attend day centres within Herefordshire but the Service also covers Service Users in neighbouring counties where the Service User is registered with a Herefordshire GP. Specialist Care Homes E.g. Stanley House, Martha Trust and Poachers Cottage, Ledbury ABI. Any acceptance and exclusion criteria and thresholds The service will be available to the population registered with a Herefordshire GP Practice. In addition the NHS across England operates informal reciprocal arrangements whereby children or adults who have temporary residence in an area other than where they are registered with their own GP can access a repair service (Section 4) e.g. when visiting relatives or on holiday.