## Sheet1
| User Requirement Specification | CPI Tender Reference & Title:- | 1590 | Framework for the Supply Delivery Installation & Commissioning of a Double Planetary Mixer
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column C, of which the total Quality element score of this sheet will equate to value in cell F6 | Scoring Criteria 0 = no response/answer not relevant to requirement 1 = limited response to requirements 2 = less than adequate, filfils some of the basic requirements 3 = adequate response, fulfils all basic requirements 4 = good response, fulfils all requirements to a high standard 5 = excellent response, fulfils requirements to an exceptionally high standard |  | The total percentage available which will contribute towards the Quality scoring element of this tender | 50%
USER REQUIREMENT |  |  |  | TENDERER RESPONSE
Requirement Codes |  |  |  | Guidance
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as desirable) |  |  | Tenderers must provide responses within the following columns; 1. Column F - please provide details of your equipment functionality in relation to CPI requirements specified in column E 2. Column G - please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response |  | Y
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  |  | N
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
# | Parameter | Requirement | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N) | N
System Performance
| Key instrument features and operation | Labotary scale double planetary mixer system, operating on principal of rotation of blades on opposing axis to generate shear. | M |  |  | 1 | 0 | 0
|  | Fully hydraulic system for raising/lowering mixing jar. | D |  |  | 1 | 0 | 1
|  | Standard control panel included. Please provide details of readouts and ability to operate system. | D |  |  | 1 | 0 | 2
|  | Instrument cabable of handling pigments and metallic particles. Please provide details. | M |  |  | 2 | 0 | 2
|  | Instrument cabable of handling nano-sized versions of the above materials. Please provide details. | D |  |  | 2 | 0 | 3
|  | Capability to heat reaction vessel. Please detail available temperature range. | M |  |  | 3 | 0 | 3
|  | Ability to cool mixing vessels during operation to limit heat build up | M |  |  | 4 | 0 | 3
|  | Bench to mount system included in the quote. | D |  |  | 4 | 0 | 4
| Safety of system | Fitted with systems to safeguard operators eg covers, limit switches, emergency stops. Please provide details. | M |  |  | 5 | 0 | 4
|  | For any compressed air requirement, please detail pressures for standard operation, and include design pressure and any relief requirements. | M |  |  | 6 | 0 | 4
|  | ATEX rated system for operation with flammable materials. Please provide appropriate documentation in support of this. | D |  |  | 6 | 0 | 5
| Vacuum system | Vacuum system with piping and seals capable of withstanding and maintaining vacuuum. | M |  |  | 7 | 0 | 5
|  | Vacuum pump for standard use included in the purchase. | M |  |  | 8 | 0 | 5
| Reaction monitoring | Preferred to have 2 x sight ports within the system for monitoring. | D |  |  | 8 | 0 | 6
|  | Reaction temperature monitoring capability with inbuilt with temperature readout. | M |  |  | 9 | 0 | 6
| Mixing vessels | Efficient sealing system to allow quick interchange of mixing vessels. | M |  |  | 10 | 0 | 6
|  | Vessel size between 250 - 600 mL. Please detail minimum and maximum volumes that can be used. | M |  |  | 11 | 0 | 6
|  | Multiple vessels included to allow interchange between formulations (Standard system to include 3 x each of ~250 mL and ~500 mL sizes). | M |  |  | 12 | 0 | 6
|  | Please provide pricing details for further mixing vessels of each size. | O |  |  | 12 | 1 | 6
|  | Please provide details of capability to use other size mixing vessels with this instrument, and any additional purchases required for this. | O |  |  | 12 | 2 | 6
| Mixing blade design | Interchangeable blade design for different formulation requirement. Please detail options and capabilities of different blade geometries. | M |  |  | 13 | 2 | 6
|  | Ability to mix at speeds from 10 to 70 rpm. | M |  |  | 14 | 2 | 6
|  | Ability to mix formulations ranging from 500 to 1 million centipoise. | M |  |  | 15 | 2 | 6
|  | Ability to mix highly viscous formulations up to 5 million centipoise. | M |  |  | 16 | 2 | 6
|  | Standard system to include 2 of each available mixing blade design | M |  |  | 17 | 2 | 6
|  | System setup to prevent accumulation of material on the sidewalls or bottom of the container. | M |  |  | 18 | 2 | 6
|  | Please provide pricing information for spare mixing blade systems. | O |  |  | 18 | 3 | 6
|  | Please provide pricing information for mixing blades coated with alternate materials eg Teflon, AlOx. | O |  |  | 18 | 4 | 6
| Spares/Consumables etc | Sealing ring systems supplied for working with different solvents eg teflon, viton, EPDM. Please provide spares and pricing for further parts as required. | M |  |  | 19 | 4 | 6
|  | To include recommended method for cleaning mixer pots and blades. Please include pricing of any required consumables as an option. | D |  |  | 19 | 4 | 7
|  | Please provide a list detailing any further consumables that may be required. | O |  |  | 19 | 5 | 7
| Additional capability | Ability to work in alternate atmospheres eg nitrogen, argon gas. Please provide details. | D |  |  | 19 | 5 | 8
|  | Additional systems to aid in sample retrieval eg Platen to push samples out of the pot. Please provide details of price and any additional considerations required eg additional bench space. | D |  |  | 19 | 5 | 9
| Installation and Utilities |  |  |  |  | 19 | 5 | 9
| Installation | Installation must be performed by qualifed engineers at CPI NetPark facility | M |  |  | 20 | 5 | 9
| Site Requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment | M |  |  | 21 | 5 | 9
| Electrical Power | Where possible, the system should operate from a standard UK mains supply (240Vac, 50Hz, single phase). Please indicate if a 3 phase supply is required. | M |  |  | 22 | 5 | 9
| Software and Connectivity |  |  |  |  | 22 | 5 | 9
| Connectivity | Please detail connectivity for data logging eg to monitor blade speed and temperature during a mixing run. | M |  |  | 23 | 5 | 9
| Regulatory, Support and Warranties |  |  |  |  | 23 | 5 | 9
| CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided | M |  |  | 24 | 5 | 9
| Training | Full training should be provided along with access to technical support | M |  |  | 25 | 5 | 9
| Warranties | Warranty period should cover at least 18 months from installation | M |  |  | 26 | 5 | 9
| Extended Warranty | Extended warranty of up to 5 years | M |  |  | 27 | 5 | 9
| Service Support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. | M |  |  | 28 | 5 | 9
|  | Details of preventative maintenance programmes required (if any) above Warranty requirements | M |  |  | 29 | 5 | 9
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com