AW12 VAR – System Technical Specification Questionnaire

| Weighting = 0 | AW12-1.1 | Please confirm you are able to meet the below minimum requirements. RF output level from -30dBm to +5dBm and 360 degree phase control RF frequency operation pulsed at 2998.5MHz 3µs 10-400Hz RF amplitude control stability of 0.1%. RF phase control stability of 0.05 degrees. Measurement of the klystron forward and reflected power level via waveguide directional coupler and measurement of forward and reflected power at the cavity structure via waveguide directional coupler. The LLRF system should be enclosed in a 19 inch rack enclosure The LLRF systems will need to be delivered and installed at Daresbury Laboratory where the agreed performance will also need to be demonstrated to the customer’s satisfaction. Delivery and installation must be before ( October 2017) |
| --- | --- | --- |
| 0 = No score Applied 1 = Less Important 2 = Important 3 = More Important | Bidder Guidance | The Bidder shall answer Yes or No Yes = Pass No = Fail |
|  | Scoring Criteria | Mandatory Pass/Fail No score Applied |
|  | Bidder Response | Mandatory Drop Down Menu – Yes/No |
| Enter Response Here |  |  |

| Weighting = 1 | AW12-1.2 | With reference to question AW12-1.1, please describe how the following requirements will be met: RF output level RF frequency operation pulsed RF amplitude control stability of 0.1%. RF phase control stability of 0.05 degrees. measurement of the klystron forward and reflected power level via waveguide directional coupler and measurement of forward and reflected power at the cavity structure via waveguide directional coupler. the LLRF system should be enclosed in a 19 inch rack enclosure Please refer as necessary to supporting documents that include sufficient description, drawings, illustration, schematics, calculated theoretical values, simulation models and results. |
| --- | --- | --- |
| 0 = No score Applied 1 = Less Important 2 = Important 3 = More Important | Bidder Guidance | Compliance Criteria |
|  | Scoring Criteria | 0 = Completely fails to meet required standard or does not provide a proposal. 20 = Proposal significantly fails to meet the standards required, contains significant shortcomings and/or is inconsistent with other proposals. 40 = Proposal falls short of achieving expected standard in a number of identifiable material respects. 60 = Proposal meets the required standard in most material respects, but is lacking or inconsistent in others. 80 = Proposal meets the required standard in all material respects. 100 = Proposal meets the required standard in all material respects and exceeds some or all of the major requirements. |
|  | Bidder Response | Multiple Line Answer (Maximum character count = 4096) |
| Enter Response Here |  |  |

| Weighting = 1 | AW12-1.3 | Please explain how your solution will be able to provide the control information to enable the customer to interface the system remotely via the CLARA control system which is an EPICS system Please refer as necessary to supporting documents that include sufficient description, drawings, illustration, schematics, calculated theoretical values, simulation models and results. |
| --- | --- | --- |
| 0 = No score Applied 1 = Less Important 2 = Important 3 = More Important | Bidder Guidance | Compliance Criteria |
|  | Scoring Criteria | 0 = Completely fails to meet required standard or does not provide a proposal. 20 = Proposal significantly fails to meet the standards required, contains significant shortcomings and/or is inconsistent with other proposals. 40 = Proposal falls short of achieving expected standard in a number of identifiable material respects. 60 = Proposal meets the required standard in most material respects, but is lacking or inconsistent in others. 80 = Proposal meets the required standard in all material respects. 100 = Proposal meets the required standard in all material respects and exceeds some or all of the major requirements. |
|  | Bidder Response | Multiple Line Answer (Maximum character count = 4096) |
| Enter Response Here |  |  |

| Weighting = 3 | AW12-1.4 | Lower Level of amplitude and phase stability The Bidder should indicate the level of performance that they guarantee Please refer as necessary to supporting documents that include sufficient description, drawings, illustration, schematics, calculated theoretical values, simulation models and results. |
| --- | --- | --- |
| 0 = No score Applied 1 = Less Important 2 = Important 3 = More Important | Bidder Guidance | Compliance Criteria |
|  | Scoring Criteria | 0 = Level of <0.05 degrees stability 20 = Minimum level of 0.05 degrees stability met 40 = 0.04 degrees stability 60 = 0.03 degrees phase stability 80 = 0.02 degrees phase stability 100 = 0.01 degrees phase stability |
|  | Bidder Response | Multiple Line Answer (Maximum character count = 4096) |
| Enter Response Here |  |  |

| Weighting = 3 | AW12.1-5 | The RF pulse structure During conditioning of cavity structure it may be necessary to reduce the RF pulse length to 0.25µs. The Bidder should indicate the pulse width that they guarantee above the minimum specified Please refer as necessary to supporting documents that include sufficient description, drawings, illustration, schematics, calculated theoretical values, simulation models and results. |
| --- | --- | --- |
| 0 = No score Applied 1 = Less Important 2 = Important 3 = More Important | Bidder Guidance | Compliance Criteria |
|  | Scoring Criteria | 0 = Operating at pulse width at >3µS 20 = Ability to operate at pulse width at 3µS 40 = Ability to operate at pulse width of 2uS 60 = Ability to operate at pulse width of 1uS 80 = Ability to operate at pulse width of 0.5uS 100 = Ability to operate at pulse width of 0.25uS |
|  | Bidder Response | Multiple Line Answer (Maximum character count = 4096) |
| Enter Response Here |  |  |

| Weighting = 1 | AW12-1.6 | Calibration function Bidders are to confirm whether their proposed system will include a calibration function that measures the phase drift of the installed probe cable so that this can be compensated for therefore removing any error due to environmental factors. Please confirm that the system will have this capability |
| --- | --- | --- |
| 0 = No score Applied 1 = Less Important 2 = Important 3 = More Important | Bidder Guidance | Yes/No Criteria |
|  | Scoring Criteria | Yes = 80 No = 0 |
|  | Bidder Response | Yes/No |
| Enter Response Here |  |  |

| Weighting = 1 | AW12-1.7 | Power Measurements Bidders are to confirm whether their proposed system will include power measurements of the two RF loads for each travelling wave structure Please confirm that the system will have this capability |
| --- | --- | --- |
| 0 = No score Applied 1 = Less Important 2 = Important 3 = More Important | Bidder Guidance | Yes/No Criteria |
|  | Scoring Criteria | Yes = 80 No = 0 |
|  | Bidder Response | Yes/No |
| Enter Response Here |  |  |