## Drop Codes (hidden)
RoomClassification | IA | QualificationTest
Not Classified | Direct Impact | Yes
Grade A | Indirect Impact | No
Grade B | No Impact | n/a
Grade C
Grade D

## Lab scale homogeniser
User Requirement Specification |  |  | CPI Internal Tender Reference & Title:- | 1064 - Lab & Pilot Scale Homogenisers
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G. | 30%
USER REQUIREMENT |  |  |  |  | Tenderer Response
Requirement codes |  |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as Desirable) |  |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  | Requirement |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
| Equipment principle | Benchtop unit for use in a R&D laboratory. The equipment will have a continuous mode of operation at high pressure for the disruption of biological cells. The equipment will enable testing at small process volume scales but provide results which are comparable to those achieved with larger scale equipment based on hydraulic piston technology. Note: This lab homogeniser is tendered with a pilot homogeniser and so must be scalable between the two. |  | M
| Application | The system should be capable of handling a variety of biological organisms including but not limited to bacterial and yeast |  | M
Technical Specification
| Operating pressure | Maximum : ≥ 1000 barg Minimum : To be confirmed by Vendor Must be configurable to allow change of pressures |  | M
| Homogeniser type | Piston based system to allow scale-up between existing systems at CPI |  | M
| Design / Dimensions | Robust and compact design suitable for a bench-top location. Vendor to confirm weight. |  | M
| Feed pump | Equipment should not require an additional feed pump |  | M
| Cooling | Ability to cool the product in-line. Vendor to provide details of cooling specifications and the utilities required. |  | M
| Minimum feed volume | Minimum feed volume of ≤5mls through the homogeniser. Vendor to detail the minimum volume allowable |  | M
| Safety | Built in Safety protection from overpressure |  | M
| Cleaning in Place | System can be Cleaned In Place (CIP) without dismantling the system. Compatible with sodium hydroxide cleaning solutions. |  | M
| Display | Local pressure, temperature displays (flowrate display optional) |  | M
| Operating flow rate | Maximum : to be confirmed by Vendor Minimum : to be confirmed by Vendor Note: See the minimum feed volume, it is proposed to process 5mls to 100's of ml for this homogeniser |  | D
| Data recording | Ability to trend pressure, temperature & flowrate via supplied computer software |  | D
| Inlet and outlet connections | Tri-clamp connections |  | D
| Configuration/type change | Option to change between technologies (i.e. piston versus microfluidics). Vendor to provide details if available |  | O
Engineering Sepcification
| Design specification | Santiary design to allow CIP and free draining appropriate of the equipment |  | M
| Material of Construction (MoC) | Preferrable stainless steel in its construction (on the product line) or alternative which is compatible with acid and base cleaning solutions |  | D
Installation & Utilities
| Installation | Installation must be performed by qualifed vendor engineers at CPI IBB facility |  | M
| Site requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment if required |  | M
| Electrical Power | The system should operate from a standard UK mains supply (240Vac, 50Hz, single phase) |  | M
| Other Utilities | The system will not require a 3 phase supply, gases or cryogenic gases |  | D
Software and Connectivity
| Connectivity | Computer Interface (e.g. ethernet, USB etc) |  | D
Regulatory, Support & Warranties
| CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided |  | M
| Training | Full training should be provided along with access to technical support |  | D
| Warranties | Warranty period should cover at least 18 months from installation |  | M
| Extended warranty | Extended warranty of up to 5 years |  | D
| Service support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. |  | D
| Service support | Details of preventative maintenance programmes required (if any) above Warranty requirements |  | D
OFFICE USE ONLY - This section below this line is for CPI Evaluation and Internal USE only
| Notes for Evaluation Team: |  | Name of Bidder to be inserted here | 1064 - Lab & Pilot Scale Homogenisers | 0
Please assess all of the answers given, with particular attention to the Mandatory answers provided, taking into account, the Desirable and Optional answers too, however, you do not need to score each individual answer, but based on the scoring range 0-5, just give one singluar score (default set at "5" in the red and yellow box below, and as we have to provide feedback to bidders, you must also provide comments on which section(s) above where you felt that they did not meet the requirements.
| Evaluation Panel |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com

## Pilot scale homogeniser
User Requirement Specification |  |  | CPI Internal Tender Reference & Title:- | 1064 - Lab & Pilot Scale Homogenisers
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 40%
USER REQUIREMENT |  |  |  |  | Tenderer Response
Requirement codes |  |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as Desirable) |  |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  | Requirement |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
| Equipment principle | Standalone unit for use in a pilot scale laboratory. The equipment will have a continuous mode of operation at high pressure for the disruption of biological cells. The equipment will enable testing at pilot process volume scales (typically 50-200L) but provide results which are comparable & scalable to those achieved with lab & larger scale homogensiers based on hydraulic piston technology. Note: This pilot homogeniser is tendered with a lab homogeniser and so must be scalable between the two. |  | M
| Application | The system should be capable of handling a variety of biological organisms including but not limited to bacterial and yeast |  | M
Technical Specification
| Operating pressure | Maximum : ≥ 1000 barg Minimum : To be confirmed by Vendor Must be configurable to allow change of pressures |  | M
| Operating flow rate | Maximum : >75 L/h Minimum : to be confirmed by Vendor |  | M
| Homogeniser type | Piston based system to allow scale-up between existing systems at CPI and the laboratory based homogensier |  | M
| Design / Dimensions | Robust and compact design, with the abiity to move the equipment easily between locations (i.e. on castors). Vendor to supply footprint |  | M
| Feed pump | Integrated feed pump for controlled flowrate |  | M
| Cooling | Ability to cool the product in-line. Vendor to provide details of cooling specifications and the utilities required. |  | M
| Minimum feed volume | Minimum feed volume of ≤1L through the homogeniser. Vendor to detail the minimum volume allowable |  | M
| Safety | Built in Safety protection from overpressure |  | M
| Cleaning in Place | System can be Cleaned In Place (CIP) without dismantling the system. Compatible with sodium hydroxide cleaning solutions. |  | M
| Display | Local pressure, flowrate and temperature displays |  | M
| Data recording | Ability to trend pressure, temperature & flowrate via supplied computer software |  | D
| Inlet and outlet connections | Tri-clamp connections |  | D
Engineering Sepcification
| Design specification | Santiary design to allow CIP and free draining appropriate of the equipment |  | M
| Material of Construction (MoC) | Preferrable stainless steel in its construction (on the product line) or alternative which is compatible with acid and base cleaning solutions |  | D
| ATEX | Vendor to specify if equipment can be ATEX Zone 2 (T2 IIA) rated. |  | O
Installation & Utilities
| Installation | Installation must be performed by qualifed vendor engineers at CPI IBB facility |  | M
| Site requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment if required |  | M
| Electrical Power | The system should operate from a standard UK mains supply (240Vac, 50Hz, single phase or 3 phase supply). Vendor to supply details |  | M
| Other Utilities | Vendor to specify the utilities required (not already specified previously) to include but not limited to instrument air pressure, water supply pressure, process air pressure |  | D
Software and Connectivity
| Connectivity | Computer Interface (e.g. ethernet, USB etc) & ability to send and receive analogue (4-20mA) signals to DCS for data reporting and for receiving instruction |  | D
Regulatory, Support & Warranties
| CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided |  | M
| Training | Full training should be provided along with access to technical support |  | D
| Warranties | Warranty period should cover at least 18 months from installation |  | M
| Extended warranty | Extended warranty of up to 5 years |  | D
| Service support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. |  | D
| Service support | Details of preventative maintenance programmes required (if any) above Warranty requirements |  | D
OFFICE USE ONLY - This section below this line is for CPI Evaluation and Internal USE only
| Notes for Evaluation Team: |  | Name of Bidder to be inserted here | 1064 - Lab & Pilot Scale Homogenisers | 0
Please assess all of the answers given, with particular attention to the Mandatory answers provided, taking into account, the Desirable and Optional answers too, however, you do not need to score each individual answer, but based on the scoring range 0-5, just give one singluar score (default set at "5" in the red and yellow box below, and as we have to provide feedback to bidders, you must also provide comments on which section(s) above where you felt that they did not meet the requirements.
| Evaluation Panel |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com