APPENDIX 1 TO SPECIAL NOTICES AND INSTRUCTIONS TO TENDERERS

TENDER EVALUATION MATRIX – JFC9a/00008

Company:

TENDER EVALUATION MATRIX

| Assessor: |  |  |  |  |  |  |  |  |  | Date: |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Serial | Criteria | Score |  |  |  |  |  | X Weighting |  | = Score | Comments |
| 1 | Essential requirements | 0 | 1 | 2 |  | 3 |  |  |  |  |  |
| 1.1 | Stainless steel (SS) test rig capable of operation at temperature 300°C and pressure of 200bar, with connecting 1/8" Swagelok SS fittings throughout. | NO | YES | N/A |  | N/A |  | 10 |  |  |  |
| 1.2 | Removable SS test column dimensions of 20-25mm id and 250-300mm long | NO | Fully compliant | N/A |  | N/A |  | 10 |  |  |  |
| 1.3 | Removable test column has ease of removal via uncoupling process from the rest of the rig (<10 minutes) | NO | YES | N/A |  | N/A |  | 6 |  |  |  |
| 1.4 | Removable test column displays ease of filling and discharge of test material which can be in the size range 0.1 – 5.0 mm diameter. | NO | YES | N/A |  | N/A |  | 2 |  |  |  |
| 1.5 | Removable test column has bottom SS support mesh capable of retaining particulate matter >20m | NO | YES | N/A |  | N/A |  | 2 |  |  |  |
| 1.6 | The HPLC type pump will be able to deliver aqueous fluids typically to cover the flow range of 15 – 250 ml/min | NO | Close to the specified range but within 30% of upper and lower values | Within the specified range |  | N/A |  | 6 |  |  |  |
| 1.7 | The HPLC Pump chosen should be relatively pulse less to allow differential pressure measurements to be made | NO | YES | N/A |  | N/A |  | 6 |  |  |  |
| 1.8 | Safety pressure relief systems must be incorporated within the test rig (supplied with safety assessment matrix for operation) and be in line with safety guidelines. | NO | YES | N/A |  | N/A |  | 10 |  |  |  |
| 1.9 | Differential pressure transducer included for measurements across the test column | NO | YES | N/A |  | N/A |  | 6 |  |  |  |
| 1.10 | An oven facility capable of operation up to 300°C is needed | NO | YES | N/A |  | N/A |  | 10 |  |  |  |
| 1.11 | A spare removable SS test column fitted with bottom mesh is required | No | YES | N/A |  | N/A |  | 10 |  |  |  |
| 1.12 | Identify the need for any valve gas bleed systems at locations throughout the rig, and incorporate if necessary | NO | Valve gas bleed system incorporated at suitable points OR absent valve gas bleed system explained and justified | N/A | N/A |  | 6 |  |  |  |  |
| 1.13 | Assembled rig with front control panel separate from adjacent oven | NO | YES | N/A | N/A |  | 6 |  |  |  |  |
| 1.14 | Fluid entry to a suitable oven at test temperature will require a sufficient heat transfer route prior to entering the test column. Exit from the test column and the oven to a suitable cooling coil facility. | NO | YES | N/A | N/A |  | 6 |  |  |  |  |
| 1.15 | Delivery, commissioning and installation of HP/HT flow rig (specify number of weeks after contract award) | NO | Within 12-14 weeks | Less than 11 weeks | N/A |  | 6 |  |  |  |  |
| 1.16 | Warranty for the HP/HT flow rig | NO | 1st year | 2nd year included | 3rd year included |  | 6 |  |  |  |  |
| 1.17 | Annual Safety Assessment for the HP/HT flow rig. To include appraisal of all safety pressure relief systems (a list of what is to be covered is to be provided) | NO | 1st year | 2nd year included | 3rd year included |  | 6 |  |  |  |  |
| 1.18 | Overall pressure to be maintained via the use of a variable back pressure regulator, and effluent sampling to be undertaken manually at a suitable location immediately after the back pressure regulator | NO | YES | N/A | N/A |  | 6 |  |  |  |  |
| 1.19 | Operating manual to be provided for the HP/HT flow test facility | NO | YES | N/A | N/A |  | 6 |  |  |  |  |

Scoring Guide:

- 3	Significantly exceeds the requirement

- 2	Exceeds the requirement

- 1	Fully meets the requirement

- 0	Does not meet the requirement

The scores are then multiplied by a weighting factor giving the final score for that tenderer against that criterion.

The weightings provided are:

- 10	high importance

- 6	average importance

- 2    low importance