## Drop Codes (hidden)
RoomClassification | IA | QualificationTest
Not Classified | Direct Impact | Yes
Grade A | Indirect Impact | No
Grade B | No Impact | n/a
Grade C
Grade D

## URS Document
User Requirement Specification |  |  | CPI Internal Tender Reference & Title:- | 1378 | Microfluidic Mixing Platform
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 40%
USER REQUIREMENT |  |  |  |  | Tenderer Response
Requirement codes |  |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as Desirable) |  |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  | Requirement |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
| System Requirements | GMP compliant |  | D | Please provide details
|  | Microfluidic Production of Droplets, Emulsions and Particles |  | M
|  | Capability to utilize different stabilization/encapsulation platforms e.g. liposomes and polymers |  | M
|  | Compatible with a range of polar and non-polar solvents |  | D | Please provide details
|  | Ability to interchange microchips of different design |  | M | Please provide details
|  | If supplying two separate pieces of equipment to satisfy tender, microchip design must be replicated in research and manufacturing equipment |  | M | Please provide evidence to support this
|  | Produces droplets / particles in the range 0.1 to 0.5 micron diameter at flow rates of 80 ml/min |  | M | Please provide evidence to support this
|  | Produces droplets/particles of <0.1micron diameter at flow rates of 80 ml/min |  | M | Please provide evidence to support this
|  | Produce droplets/particles with near monomodal size distribution |  | M | Please provide evidence to support this
|  | Pulsless flow |  | D
|  | Continuous flow pumping system, If supplying two separate pieces of equipment to satisfy tender, research equipment must be capable of delivering 20ml/min and manufacturing equipment must be capable of delivering flow rates of 80ml/min (producing material at 0.1-0.5um diameter) |  | M | Please provide evidence to support this
|  | Continuous flow pumping system, If supplying a single piece of equipment to satisfy tender, must be able to deliver flow rates of 80ml/min (producing material at 0.1-0.5um diameter) |  | M | Please provide evidence to support this
|  | Continuous flow capable of delivering batch sizes of 10L |  | M | Please supply details
|  | Must supply cartridges suitable for multiple use |  | M | Provide details regarding cost of cartridges for both research and manufacturing equipment and provide number of runs per cartridge
|  | Must include computer |  | M | Provide costing
Installation & Utilities
| Installation | Installation must be performed by qualifed engineers at CPI NetPark facility |  | M
| Site requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment |  | M
| Electrical Power | The system should operate from a standard UK mains supply (240Vac, 50Hz, single phase) |  | M
| Other Utilities | The system should be bench mounted |  | D
| Other Utilities | The system will not require a 3 phase supply, cooling water, gases or cryogenic gases |  | D
Software and Connectivity
| Connectivity | Computer Interface (e.g. ethernet, USB etc) |  | D
Regulatory, Support & Warranties
| CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided |  | M
| Training | Full training should be provided along with access to technical support |  | D
| Warranties | Warranty period should cover at least 12 months from installation |  | M
| Extended warranty | Extended warranty of up to 5 years |  | D
| Service support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. |  | D
| Service support | Details of preventative maintenance programmes required (if any) above Warranty requirements |  | D
OFFICE USE ONLY - This section below this line is for CPI Evaluation and Internal USE only
| Notes for Evaluation Team: |  | Name of Bidder to be inserted here | 1378 | Microfluidic Mixing Platform
Please assess all of the answers given, with particular attention to the Mandatory answers provided, taking into account, the Desirable and Optional answers too, however, you do not need to score each individual answer, but based on the scoring range 0-5, just give one singluar score (default set at "5" in the red and yellow box below, and as we have to provide feedback to bidders, you must also provide comments on which section(s) above where you felt that they did not meet the requirements.
| Evaluation Panel |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com