## Sheet1
| User Requirement Specification | CPI Tender Reference & Title:- | 1590 | Framework for the Supply Delivery Installation & Commissioning of a Double Planetary Mixer
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column C, of which the total Quality element score of this sheet will equate to value in cell F6 | Scoring Criteria 0 = no response/answer not relevant to requirement 1 = limited response to requirements 2 = less than adequate, filfils some of the basic requirements 3 = adequate response, fulfils all basic requirements 4 = good response, fulfils all requirements to a high standard 5 = excellent response, fulfils requirements to an exceptionally high standard |  | The total percentage available which will contribute towards the Quality scoring element of this tender | 50%
USER REQUIREMENT |  |  |  | TENDERER RESPONSE
Requirement Codes |  |  |  | Guidance
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as desirable) |  |  | Tenderers must provide responses within the following columns; 1. Column F - please provide details of your equipment functionality in relation to CPI requirements specified in column E 2. Column G - please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response |  | Y
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  |  | N
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
# | Parameter | Requirement | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N) | N
System Performance
M1 | Key instrument features and operation | Laboratory scale double planetary mixer system suitable for efficient mixing of viscous mixtures. | M |  |  | 1 | 0 | 0
M2 |  | Instrument cabable of handling pigments and metallic particles | M |  |  | 2 | 0 | 0
M3 |  | Materials of construction and sealing ring systems suitable for working with solvents. | M |  |  | 3 | 0 | 0
M4 |  | Capability to heat reaction vessel. Please detail available temperature range. | M |  |  | 4 | 0 | 0
M5 |  | Ability to cool mixing vessels during operation to limit heat build up | M |  |  | 5 | 0 | 0
D1 |  | Inbuilt reaction temperature monitoring capability with temperature readout. | D |  |  | 5 | 0 | 1
M6 | Safety of system | Fitted with systems to safeguard operators eg covers, emergency stops. | M |  |  | 6 | 0 | 1
M7 |  | Please detail extraction requirements for use of system with volatile solvents. | M |  |  | 7 | 0 | 1
O1 |  | Please provide detail of additional costs required to provide ATEX rated system | O |  |  | 7 | 1 | 1
M8 | Vacuum system | System capable of operating under vacuum, with piping and seals capable of withstanding and maintaining vacuuum. | M |  |  | 8 | 1 | 1
M9 |  | Vacuum pump for standard use included in the purchase. | M |  |  | 9 | 1 | 1
M10 | Mixing vessels | Interchangeable mixing vessels to allow for efficient workflows for multiple samples. | M |  |  | 10 | 1 | 1
D2 |  | Multiple vessels included in purchase to allow interchange between formulations | D |  |  | 10 | 1 | 2
M11 |  | Volume of mixing chamber less than or equal to 1 L. Please detail mixing chamber volume and minimum and maximum volumes that can be used. | M |  |  | 11 | 1 | 2
D3 |  | Ability to replace O-rings and sealing systems to use alternative solvents. Please provide details. | D |  |  | 11 | 1 | 3
M12 | Mixing blade design | Interchangeable blade design for different formulation requirement and ease of replacement. Please detail options and capabilities of different blade geometries. | M |  |  | 12 | 1 | 3
M13 |  | Blades able to operate at varying speeds up to 1000 rpm. | M |  |  | 13 | 1 | 3
M14 |  | Ability to mix formulations with viscosity ranging from 10 to 2500 Pa.s | M |  |  | 14 | 1 | 3
O2 |  | Ability to mix highly viscous formulations with viscosity up to 10000 Pa.s | O |  |  | 14 | 2 | 3
M15 |  | System equipped with scrapers to prevent accumulation of material on the sidewalls or bottom of the container. | M |  |  | 15 | 2 | 3
O3 | Spares/Consumables etc | Please provide pricing information for spare mixing blade systems. | O |  |  | 15 | 3 | 3
O4 |  | Please provide pricing details for further mixing vessels of each size. | O |  |  | 15 | 4 | 3
O5 |  | Please provide details of capability to use other size mixing vessels with this instrument, and any additional purchases required for this. | O |  |  | 15 | 5 | 3
D4 | Additional capability | Additional systems to aid in sample retrieval eg Platen to push samples out of the pot. Please provide details of price and any additional considerations required eg additional bench space. | D |  |  | 15 | 5 | 4
0 | Installation and Utilities |  |  |  |  | 15 | 5 | 4
M16 | Installation | Installation must be performed by qualifed engineers at CPI NetPark facility | M |  |  | 16 | 5 | 4
M17 | Site Requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment | M |  |  | 17 | 5 | 4
M18 | Electrical Power | Where possible, the system should operate from a standard UK mains supply (240Vac, 50Hz, single phase). Please indicate if a 3 phase supply is required. | M |  |  | 18 | 5 | 4
0 | Software and Connectivity |  |  |  |  | 18 | 5 | 4
M19 | Connectivity | Please detail connectivity for data logging eg to monitor blade speed and temperature during a mixing run. | M |  |  | 19 | 5 | 4
0 | Regulatory, Support and Warranties |  |  |  |  | 19 | 5 | 4
M20 | CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided | M |  |  | 20 | 5 | 4
M21 | Training | Full training should be provided along with access to technical support | M |  |  | 21 | 5 | 4
M22 | Warranties | Warranty period should cover at least 18 months from installation | M |  |  | 22 | 5 | 4
D5 | Extended Warranty | Extended warranty of up to 5 years | D |  |  | 22 | 5 | 5
M23 | Service Support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. | M |  |  | 23 | 5 | 5
M24 |  | Details of preventative maintenance programmes required (if any) above Warranty requirements | M |  |  | 24 | 5 | 5
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com