## Sheet1
Guidance for completion - When inserting additional rows in the requirements table below (Starting at row 16), please ensure that you drag the formulas from Column A down from the row above into the newly inserted row. - Do not delete the top and bottom rows of the requirements table (Starting Row 16 and finishing Row 48). Only insert rows between these points and do not add any lines after these points. - Hide row 5 if lots arent applicable
| User Requirement Specification | CPI Tender Reference & Title:- | 1755 | Framework for the Supply Delivery Installation & Commissioning of Fatigue Test Equipment
Lot # and Title:-
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column C, of which the total Quality element score of this sheet will equate to value in cell F6 | Scoring Criteria 0 = no response/answer not relevant to requirement 1 = limited response to requirements 2 = less than adequate, filfils some of the basic requirements 3 = adequate response, fulfils all basic requirements 4 = good response, fulfils all requirements to a high standard 5 = excellent response, fulfils requirements to an exceptionally high standard |  | The total percentage available which will contribute towards the Quality scoring element of this tender | 50%
USER REQUIREMENT |  |  |  | TENDERER RESPONSE
Requirement Codes |  |  |  | Guidance
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as desirable) |  |  | Tenderers must provide responses within the following columns; 1. Column F - please provide details of your equipment functionality in relation to CPI requirements specified in column E 2. Column G - please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response |  | Y
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  |  |  | N
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
# | Parameter | Requirement | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N) | N
System Performance
M1 | Equipment | Capable of dynamic testing at +/- 10 kN | M |  |  | 1 | 0 | 0
M2 |  | Environmental chamber capable of sub-ambient operation using liquid nitrogen (down to -70°C or lower) and heating to at least 200 °C | M |  |  | 2 | 0 | 0
M3 |  | Capable of dynamic testing in tension, bending and compression with all necessary fittings and grips supplied for use within and outside the chamber | M |  |  | 3 | 0 | 0
O1 | Equipment | Capable of static testing as well | O |  |  | 3 | 1 | 0
0 | Installation and Utilities |  |  |  |  | 3 | 1 | 0
M4 | Installation | Installation must be performed by qualifed engineers at CPI NetPark facility | M |  |  | 4 | 1 | 0
M5 | Site Requirements | A full site requirements guide must be provided as part of the Tender response, including requirements for specialist lifting and manual handling equipment | M |  |  | 5 | 1 | 0
M6 | Electrical Power | The system should operate from a standard UK mains supply (240Vac, 50Hz, single phase) | M |  |  | 6 | 1 | 0
M7 | Other Utilities | The system should not require any utilities apart from mains supply (an exception would be occasional need for liquid nitrogen for sub-ambient use of the chamber) | M |  |  | 7 | 1 | 0
0 | Software and Connectivity |  |  |  |  | 7 | 1 | 0
M8 | Connectivity | Computer Interface (e.g. ethernet, USB etc). Supplier to liase with CPI IT to overcome any interference in equipment software from CPI intranet features such as firewalls, Electronic Lab Notebook software etc. | M |  |  | 8 | 1 | 0
0 | Regulatory, Support and Warranties |  |  |  |  | 8 | 1 | 0
M9 | CE Marking | The Equipment shall be permanently CE Marked to confirm that it is in accordance with all relevant EU Directives, and a suitable conformance certificate shall be provided | M |  |  | 9 | 1 | 0
M10 | Training | Full training should be provided along with access to technical support | M |  |  | 10 | 1 | 0
M11 | Warranties | Warranty period should cover at least 18 months from installation | M |  |  | 11 | 1 | 0
M12 | Extended Warranty | Extended warranty of up to 5 years | M |  |  | 12 | 1 | 0
M13 | Service Support | On site support within 24 hours of requesting assistance. Please provide details of expected call-out charges. | M |  |  | 13 | 1 | 0
M14 |  | Details of preventative maintenance programmes required (if any) above Warranty requirements | M |  |  | 14 | 1 | 0
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com