## Drop Codes (hidden)
RoomClassification | IA | QualificationTest
Not Classified | Direct Impact | Yes
Grade A | Indirect Impact | No
Grade B | No Impact | n/a
Grade C
Grade D

## URS Document
User Requirement Specification |  | CPI Internal Tender Reference & Title:- | 1142 | Floor Standing Lab Centrifuge and Rotors
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 60%
USER REQUIREMENT |  |  |  | Tenderer Response
Requirement codes |  |  |  | Guidance: Tenderers must provide responses within the following Columns;
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as Desirable) |  |  | 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase |  |  | 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
| Specification | A floor standing centrifuge | M
|  | Capability for centrifuging 6 x 1L samples at 8,500 RPM and 15,800 x g | M
|  | Capability for centrifuging 8 x 50mL samples at up to 20,000 RPM and up to 50,000 x g | M
|  | Rotors designed to accommodate the sample ranges defined above must be included. To avoid issues of manual handling large heavy rotors tenderer's are requested to supply rotors constructed of lightweight materials. | M
|  | Refrigeration to enable centrifugation below ambient temperature. | M
|  | Automatic recognition of rotors when installed. | M
|  | Ability to input, store and recall frequently used programs. | M
|  | Safety systems that prevent door opening until rotor has stopped. | M
|  | Ability to set acceleration and deceleration rates. | M
Installation & Utilities
| Installation | The Tenderer is required to submit weights and dimensions of all packaged equipment prior to shipment. The maximum available space to accommodate the centrifuge is 880mm (width). The centrifuge must fit into this space allowing room either side for ventilation/manoeuvrability. | M
| Site requirements | The Tenderer is required to submit weights and dimensions of all packaged equipment prior to shipment. The maximum available space to accommodate the centrifuge is 880mm (width). The centrifuge must fit into this space allowing room either side for ventilation/manoeuvrability. | M
| Electrical Power | The Tenderer shall submit detailed requirements for electrical power connections. It should be noted that a 32A 3-phase and neutral power supply is available at the intended installation location and there is a preference to make use of this for connection of the new centrifuge. | M
| Other Utilities | CPI will provide electrical power, as identified by the Tenderer, and will provide resources to connect the equipment to these services. | D
Software and Connectivity
| Connectivity |  | D
Regulatory, Support & Warranties
| CE Marking | All equipment is to be designed and manufactured to current European and UK legislation and directives as applicable and comply with the applicable EC directives in accordance with The Supply Of Machinery (Safety) Regulations 1992 (As Amended), including the CE mark, together with the supply of the required Certificate of Conformity Equipment must be CE marked including all control panels and systems. | M
| Training | The Tenderer will be responsible for supply, delivery, provision of suitable personnel for technical supervision of the positioning and connection of the equipment, commissioning and training of users. | M
| Warranties | All equipment supplied by the Tenderer shall include a minimum warranty period of 12 months, as detailed in CPIs Terms and Conditions for the supply of Goods and On-site Services attached to this document, from commissioning. | M
| Extended warranty | Options to extend the warranty period to 3 years should be included in the response to tender. | M
| Service support | The tenderer shall supply details of an ongoing maintenance contract to cover 2 years maintenance. Please provide a fixed fee for this. | D
OFFICE USE ONLY - This section below this line is for CPI Evaluation and Internal USE only
| Notes for Evaluation Team: | Name of Bidder to be inserted here | 1142 | Floor Standing Lab Centrifuge and Rotors
Please assess all of the answers given, with particular attention to the Mandatory answers provided, taking into account, the Desirable and Optional answers too, however, you do not need to score each individual answer, but based on the scoring range 0-5, just give one singluar score (default set at "5" in the red and yellow box below, and as we have to provide feedback to bidders, you must also provide comments on which section(s) above where you felt that they did not meet the requirements.
| Evaluation Panel | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com