## Drop Codes (hidden)
RoomClassification | IA | QualificationTest
Not Classified | Direct Impact | Yes
Grade A | Indirect Impact | No
Grade B | No Impact | n/a
Grade C
Grade D

## Appendix 1.1
User Requirement Specification |  |  | CPI Internal Tender Reference & Title:- | 1464 | Lot 1 - Raman Spectrometer
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 50%
USER REQUIREMENT |  |  |  |  | Tenderer Response
Requirement codes |  |  |  |  | Guidance: Tenderers must provide responses within the following Columns; 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as Desirable)
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  | Requirement |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
M1 | System Fundementals | The equipment must a Raman Spectrometer appropriate for use in the analysis of biologically-derived liquid samples. For example, in the analysis of relevant metabolite concentrations in cell culture medium and the determination of antibody quailty/aggregation in harvested bioreactor material. |  | M
M2 | Range of Operation | The system must be able to function at multiple incident wavelengths. This may be achieved either by supplying the equipment with light sources covering an adequate wavelength range for use in the analysis of biological samples or having the equipment accomodate easily interchangable laser sources. |  | M
M3 | Sample Preparation | The equipment must be capable of analysing in-process cell culture samples with little to no need for pre-purification. |  | M
D1 | Sample Format | The ability to take samples in multiple formats (cuvettes, multi-well plates, etc) is desirable. |  | D
M4 | Spectral Range | Details of the spectral range over which the system can function must be provided. |  | M
D2 | Consumables | The ability to use of generic consumables (sample containers, etc.) is prefered. |  | D
D3 | Detector | Ideally, the supplied detector should function at room temperature (with no need for cryogenic cooling). |  | D
M5 | Software | A software license agreement must be included in the provided costs. |  | M
D4 |  | The ability to analyse and store basic component data for reference in the future evaluation of more complex mixtures (as part of an automated detection and identification process), is preferred. |  | D
D5 |  | The ability to perform basic analysis of spectral data within the built-in software if desirable. |  | D
D6 | Connectivity/Compatibility | The ability for the system to be integrated with a Sartorius Ambr-250 or Ambr-15 high-throughput culture system is prefered. This will require the system to be supplied with a fibre probe. |  | D
D7 | Evidence of suitability | If possible, please provide literature demonstrating the use of the proposed piece of equipment in the area of cell culture sample analysis, bioreactor process control or antibody quality/titre assessent. |  | D
Installation & Utilities
M6 | Installation | All costs associated with delivery and installation at 1 Union Square, Darlington should be included. |  | M
M7 | Site requirements | A full site requirements guide must be provided as part of the Tender response, including a list of relevant utilities (electricity, gases, etc.), equipment dimensions and any need for specialist lifting and manual handling equipment. |  | M
M8 | Electrical Power | Requirements should be provided. |  | M
Software and Connectivity
M9 | Connectivity | Data must be easily exported through the use of multiple file formats, via either USB or network link (wired or wireless). |  | M
Regulatory, Support & Warranties
M10 | Training | Training should be provided along with access to technical support. Training should cover routine operation and troubleshooting. |  | M
M11 | Warranties | The included warranty period should cover at least 12 months from installation. |  | M
D8 | Extended warranty | The opportunity to choose an extended warranty of up to 5 years would be prefered. |  | D
M12 | Service support | Maintenance contract options should be provided (if applicable), along with all associated costs. Please also provide details of expected call-out charges (if applicable). |  | M
OFFICE USE ONLY - This section below this line is for CPI Evaluation and Internal USE only
| Notes for Evaluation Team: |  | Name of Bidder to be inserted here | 1464 | Lot 1 - Raman Spectrometer
Please assess all of the answers given, with particular attention to the Mandatory answers provided, taking into account, the Desirable and Optional answers too, however, you do not need to score each individual answer, but based on the scoring range 0-5, just give one singluar score (default set at "5" in the red and yellow box below, and as we have to provide feedback to bidders, you must also provide comments on which section(s) above where you felt that they did not meet the requirements.
| Evaluation Panel |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com

## Appendix 1.2
User Requirement Specification |  |  | CPI Internal Tender Reference & Title:- | 1464 | Lot 2 - Raman Microscope
| The evalution team will consider bidders combined responses below and give a single overall score based on the adjacent rating in column D, of which the total Quality element score of this sheet will equate to value in column G |  |  |  | The total percentage available which will contribute towards the Quality scoring element of this tender is contained within column G, of which the evaluation team will assess in the "Office Use Only" section below at the base of this form | 50%
USER REQUIREMENT |  |  |  |  | Tenderer Response
Requirement codes |  |  |  |  | Guidance: Tenderers must provide responses within the following Columns; 1. Column F - Please provide details of your equipment functionality in relation to CPI requirements specified in column E 2. Column G - Please confirm if your product complies with CPI's functionality requirements specified in Column E, by providing a Yes (Y) or No (N) response
M | Mandatory. (This must be the minimum requirement for the usage for the equipment, any further performance requirements above the minimum, must be marked as Desirable)
O | Optional requirement which may be quoted for and an assessment made by CPI prior to purchase
D | Desirable, but not necessary - if this can be supplied. Tenderer should suggest alternative if appropriate
Parameter |  | Requirement |  | Requirement (M/O/D) -see codes | Answer or Details with regards to meeting the Requirement | Complies? (Y/N)
System performance
M1 | System Fundementals | The equipment must a Raman Microscope appropriate for use in the analysis of both live and fixed cell samples (microbial, fungal and mammalian) alongside biologically-derived liquid samples. In regards to the liquid samples it is important that the system offer the ability to investigate things such as relevant metabolite concentrations in samples of cell culture medium and antibody quailty/aggregation in harvested bioreactor material. |  | M
M2 | Imaging | The system must be capable of imaging individual cells in both raman-mode and as a high-quality conventional microscope. |  | M
M3 | Range of Operation | The system must be able to function at multiple incident wavelengths. This may be achieved either by supplying the equipment with light sources covering an adequate wavelength range for use in the analysis of biological samples or having the equipment accomodate easily interchangable laser sources. |  | M
M4 | Sample Preparation | The equipment must be capable of analysing in-process cell culture samples with little to no need for pre-purification. |  | M
M5 | Sample Format | The ability to take samples in both microscope compatible formats and standard bulk-liquid formats (cuvettes, etc) is required. |  | M
M6 | Spectral Range | Details of the spectral range over which the system can function must be provided. |  | M
D1 | Consumables | The ability to use of generic consumables (sample containers, etc.) is prefered. |  | D
D2 | Detector | Ideally, the supplied detector should function at room temperature (with no need for cryogenic cooling). |  | D
M7 | Software | A software license agreement must be included in the provided costs. |  | M
D3 |  | The ability to analyse and store basic component data for reference in the future evaluation of more complex mixtures (as part of an automated detection and identification process), is preferred. |  | D
D4 |  | The ability to perform basic analysis of spectral data within the built-in software if desirable. |  | D
D5 | Connectivity/Compatibility | The ability for the system to be integrated with a Sartorius Ambr-250 or Ambr-15 high-throughput culture system is prefered. This will require the system to be supplied with a fibre probe. |  | D
D6 | Evidence of suitability | If possible, please provide literature demonstrating the use of the proposed piece of equipment in the area of cell culture sample analysis, bioreactor process control or antibody quality/titre assessent. |  | D
Installation & Utilities
M8 | Installation | All costs associated with delivery and installation at 1 Union Square, Darlington should be included. |  | M
M9 | Site requirements | A full site requirements guide must be provided as part of the Tender response, including a list of relevant utilities (electricity, gases, etc.), equipment dimensions and any need for specialist lifting and manual handling equipment. |  | M
M10 | Electrical Power | Requirements should be provided. |  | M
Software and Connectivity
M11 | Connectivity | Data must be easily exported through the use of multiple file formats, via either USB or network link (wired or wireless). |  | M
Regulatory, Support & Warranties
M12 | Training | Training should be provided along with access to technical support. Training should cover routine operation and troubleshooting. |  | M
M13 | Warranties | The included warranty period should cover at least 12 months from installation. |  | M
D7 | Extended warranty | The opportunity to choose an extended warranty of up to 5 years would be prefered. |  | D
M14 | Service support | Maintenance contract options should be provided (if applicable), along with all associated costs. Please also provide details of expected call-out charges (if applicable). |  | M
OFFICE USE ONLY - This section below this line is for CPI Evaluation and Internal USE only
| Notes for Evaluation Team: |  | Name of Bidder to be inserted here | 1464 | Lot 2 - Raman Microscope
Please assess all of the answers given, with particular attention to the Mandatory answers provided, taking into account, the Desirable and Optional answers too, however, you do not need to score each individual answer, but based on the scoring range 0-5, just give one singluar score (default set at "5" in the red and yellow box below, and as we have to provide feedback to bidders, you must also provide comments on which section(s) above where you felt that they did not meet the requirements.
| Evaluation Panel |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
|  |  | Name of CPI Evaluator to insert here | dd/mm/yyyy | Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements | 5
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Comments on which section(s) they felt the Bidder did not meet fully or partially the above requirements
Head Office: CPI, Wilton Centre, Wilton, Redcar, Cleveland, TS10 4RF, UNITED KINGDOM
Tel: +44 (0)1642 455 340 Email: info@uk-cpi.com Web: www.uk-cpi.com